## JEOL Info 1200 (hidden)
No. | Item | Usage (How to use) | Setup | Consumption
|  |  |  | Detail | Calculations
1 | Reagent Probe Wash1 | Alkaline Detergent for preventing contamination of RPP and RRV (Straight) | RTT-1, 2 (No.42) | RTT-1, 2: * 22ul for each aspiration (15) * This aspirated detergent will be dispensed to cuvette with 198ul (135) of pure water when System makes it to prevent the contaminations. | RTT-1: 22ul/aspiration (15) RTT-2: 22ul/aspiration (15)
2 | Reagent Probe Wash2 | Acid Detergent for preventing contamination of RPP and RRV (Straight) | RTT-1, 2 (No.43) | RTT-1, 2: * 88ul for each aspiration (60) * This aspirated detergent will be dispensed to cuvette with 132ul (90)of pure water when System makes it to prevent the contaminations. | RTT-1: 88ul/aspiration (60) RTT-2: 88ul/aspiration(60)
3 | Reagent Probe Wash3 | Detergent for Weekly Wash (5% Dilution) | CTT RTT-1, 2 | CTT: * 1.25ul(=25ul*5%) for each aspiration * 25ul of 5% diluted solution are aspirated. * 300ul(=1.25ul*240cuvettes) for each WASH2 RTT-1, 2: * 2.0ul(=40ul*5%) for each aspiration * 40ul of 5% diluted solution are aspirated. * 480ul(=2.0ul*240cuvettes) for eash WASH2 | CTT : 300ul/WASH2 RTT-1: 480ul/WASH2 RTT-2: 480ul/WASH2
4 | Cuvette Wash (Contined to next page) | Alkaline Detergent for RRV Cuvette Wash on WUD (Straight) | 2L Bottle | * 76.6ul to 90.0ul for each cycle * General Analysis: 4.5sec./cycle 76.6ul*800T/H = 61.28ml/H 90.0ul*800T/H = 72.00ml/H 61.28ml to 72.00ml for 1 hour Analysis * WASH1 (1cycle): 4.5sec./cycle 76.6ul*39min*60sec./4.5sec. = 39.832ml 90.0ul*39min*60sec./4.5sec. = 46.800ml * WASH2 (2cycle): 4.5sec./cycle 76.6ul*57min*60sec./4.5sec. = 58.216ml 90.0ul*57min*60sec./4.5sec. = 68.400ml * Cell Blank: 3.0sec./cycle 76.6ul*17min*60sec./3.0sec. = 26.044ml 90.0ul*17min*60sec./3.0sec. = 30.600ml | 76.6 - 90.0ul/cycle *General Analysis: 61.28 - 72.00 ml/H *WASH1 (1cycle): 39.832 - 46.800 ml/WASH1 *WASH2 (2cycle): 58.216 - 68.400 ml/WASH2 *Cell Blank: 26.044 - 30.600 ml/CB
| Cuvette Wash | Detergent for daily Wash (10% Dilution) | CTT RTT-1, 2 (No.44) | CTT: * 2.5ul(=25ul*10%) for each aspiration * 25ul of 5% diluted solution are aspirated. * 600ul(=2.5ul*240cuvettes) for each WASH2 RTT-1, 2: * 4.0ul(=40ul*10%) for each aspiration * 40ul of 10% diluted solution are aspirated and will be dispensed to cuvette with 40ul of pure water. * 960ul(=4.0ul*240cuvettes) for each WASH2 | CTT : 600ul/WASH2 RTT-1: 960ul/WASH2 RTT-2: 960ul/WASH2
5 | Cuvette Conditioner | Conditioner for RRV Cuvette Wash on WUD (Straight) | 2L Bottle | * 30.0ul to 46.7ul for each cycle * General Analysis: 4.5sec./cycle 30.0ul*800T/H = 24.00ml/H 46.7ul*800T/H = 38.08ml/H 61.28ml to 72.00ml for 1 hour Analysis * WASH1 (1cycle): 4.5sec./cycle 30.0ul*39min*60sec./4.5sec. = 15.600ml 46.7ul*39min*60sec./4.5sec. = 24.284ml * WASH2 (2cycle): 4.5sec./cycle 30.0ul*57min*60sec./4.5sec. = 22.800ml 46.7ul*57min*60sec./4.5sec. = 35.492ml * Cell Blank: 3.0sec./cycle 30.0ul*17min*60sec./3.0sec. = 10.200ml 46.7ul*17min*60sec./3.0sec. = 15.878ml | 30.0 - 46.7ul/cycle *General Analysis: 24.0 - 38.08 ml/H *WASH1 (1cycle): 15.600 - 24.284 ml/WASH1 *WASH2 (2cycle): 22.800 - 35.492 ml/WASH2 *Cell Blank: 10.200 - 15.878 ml/CB
6 | Incubation Bath Oil | Circulating Oil for Reaction Bath (Straight) | 2L Bottle | * Baseline: under 100ml for each day * Average : 40ml for each day | Baseline: under 100ml/day Average : 40ml/day
7 | Lamp Coolant | Circulating Anticorrosion Liquid for Cooling of Detector Lamp (5% Dilution) | Lamp-cooling Liquid Tank | * This depends on environment of installed place. | ---
8 | ISE Buffer | Buffer of ISE (Straight) | 2L Bottle | In case of Continuous Analysis (Waiting time: 0 sec. to less than 5 sec.): * 3.00ml for each Serum sample * 4.50ml for each Urine Sample In case of Discontinuous Analysis: It depends on waiting time from last analysis. * Waiting time: 5 sec. to less than 3 min. Prewash 2times(1.50ml) + Analysis: Serum: 4.50ml Urine: 6.00ml * Waiting time: 3 min. to less than 10 min. Prewash 3times(2.25ml) + Analysis Serum: 5.25ml Urine: 6.75ml * Waiting time: 10 min. and over Prewash 4times(3.00ml) + Analysis Serum: 6.00ml Urine: 7.50ml In case of Wash Electrode and WASH: * 4.50ml for each Wash Electrode In case of ISE line wash: * 7.50ml for each ISE line wash | Continuous Analysis (0 sec. - 5 sec.): * 3.00ml/Serum Sample * 4.50ml/Urine Sample Discontinuous Serum Analysis: * Waiting time: 5 sec. - 3 min.: 4.50ml 3 min. - 10 min.: 5.25ml 10 min. - : 6.00ml Discontinuous Urine Analysis: * Waiting time: 5 sec. - 3 min.: 6.00ml 3 min. - 10 min.: 6.75ml 10 min. - : 7.50ml Wash Electrode and WASH: * 4.50ml/Wash Electrode ISE line wash: * 7.50ml/ISE line wash
9 | ISE Detergent Solution | Detergent for Wash Electrode and WASH (Straight) | CTT | * 26ul for each aspiration * 52ul(=26ul*2 times) for each Wash Electrode | 52ul/Wash Electrode and WASH
10 | ISE Serum Standard H | Standard for ISE Calibration (Neat) | CTT | * 26ul for each aspiration * One Calibration makes 3 times to 8 times aspirations. 3 times: 78ul (=26*3) 8 times: 208ul (=26*8) | 78 - 208ul/Calibration
11 | ISE Serum Standard L | Standard for ISE Calibration (Neat) | CTT | * 26ul for each aspiration * One Calibration makes 3 times to 8 times aspirations. 3 times: 78ul (=26*3) 8 times: 208ul (=26*8) | 78 - 208ul/Calibration
12 | ISE Urine Standard H | Standard for ISE Calibration (Neat) | CTT | * 26ul for each aspiration * One Calibration makes 3 times to 8 times aspirations. 3 times: 78ul (=26*3) 8 times: 208ul (=26*8) | 78 - 208ul/Calibration
13 | ISE Urine Standard L | Standard for ISE Calibration (Neat) | CTT | * 26ul for each aspiration * One Calibration makes 3 times to 8 times aspirations. 3 times: 78ul (=26*3) 8 times: 208ul (=26*8) | 78 - 208ul/Calibration
14 | ISE Check Solution | Selectivity-check solution for Na, K (Neat) | CTT | --- | ---

## Customer Parameters 1200 (hidden)
Customer Parameters | Value |  | Ancillary | mL/ Year | Bottles/Year | Pkg/Y | Packaging | Dead Vol | SPC/ Package | SPC/bottle | Cost/year | $$ / Y | Consumo
Tests per Year (patients, incl ISE) | #REF! |  | Probe Wash 1 | 525.6 | 2.5184 | 0.5 | 5x250mL | 2mL in 40mL wedge | 24.15 | 4.83 | 12.1639 | 12.075 | 1
% dilns & reruns | 0 |  | Probe Wash 2 | #REF! | #REF! | 1.5 | 5x250mL | 2mL in 40mL wedge | 33.6 | 6.72 | #REF! | 50.4 | #REF!
Total Chem Tests/ year | #REF! |  | Probe Wash 3 | 74.048 | 1 | 0.35 | 1x500mL | 2mL in 40mL wedge | 15.75 | 15.75 | 15.75 | 5.5125 | 1
Mg Tests per year | #REF! |  | Cuvette Wash | #REF! | #REF! | 50 | 1x2L | top off on-system bottle | 12.97 | 12.97 | #REF! | 648.5 | #REF!
-or- %Mg Tests | #REF! |  | Cuvette Conditioner | #REF! | #REF! |  | 1x1L | top off on-system bottle | 18.95 | 18.95 | #REF! | 0 | #REF!
ISE tests per year | #REF! | #REF! | ISE Buffer | #REF! | #REF! |  | 1x2L | 250mL | 28.14 | 28.14 | #REF! | 0 | #REF!
serum ISE panel's per year | #REF!
urine ISE panel's per year | 0 |  | Incubation Bath Oil | 14600 | 5.0345 |  | 1x2.9L | top off on-system bottle | 279 | 279 | 1404.6207 | 0 | 6
Shifts per Day | 2 |  | Lamp Coolant | 0 | 1 |  | ~50mL |  | 23.1 | 23.1 | 23.1 | 0 | 1
Runs per Day | 12
Idle Hours per Day | 8 |  | ISE Detergent | 182.5 | 1.825 | 1 | 2x100mL |  | 24.15 | 12.075 | 22.0369 | 24.15 | 1
|  |  | ISE Cal Serum Low | 365 | 3.65 |  | 1x100mL
|  |  | ISE Cal Serum High | 365 | 3.65 |  | 1x100mL
|  |  | ISE Cal Urine Low | 365 | 3.65 |  | 1x100mL
|  |  | ISE Cal Urine High | 365 | 3.65 |  | 1x100mL
TOTAL COST | [10 empty] | #REF!
Calculated cells (grey)
Enter usage data (yellow)
Assumes ISE cal once per shift
Bottles/year calculation includes dead volume
ADDITIONAL CONSUMABLES (Generally fixed for ADVIA 1200 and are not dependent on test volume)
# Used/ Year
Cuvettes |  |  |  |  | 22 |  | 1 segment |  |  | 82 | 1804 |  | 22
Lamp |  |  |  |  | 1.716 |  | each |  |  | 77 | 132.132 |  | 2
Electrodes (assume change every 3 months) |  |  |  |  | 12 |  |  |  |  | 290 | 3480 |  | 12
Ref Electrode |  |  |  |  | 1 |  |  |  |  | 676.36 | 676.36 |  | 1
TOTAL | [10 empty] | 6092.492
Note on I5: Bayer Corporation: weekly change of wedge

## Ancill. Usage 1200 (hidden)
No. | Item | Usage (How to use) | Setup | Consumption |  | uL per Chem | uL per ISE | uL per Mg | uL per | uL per | uL per | mL Yearly use | mL Total
|  |  |  | Detail | Calculations | TEST | TEST | TEST | RUN | DAY | WEEK
1 | Reagent Probe Wash1 | Alkaline Detergent for preventing contamination of RPP and RRV (Straight) Contamination Settings | RTT-1, 2 (No.42) | RTT-1, 2: [Clean 1 settings] * 22ul for each aspiration (15) * This aspirated detergent will be dispensed to cuvette with 198ul (135) of pure water when System makes it to prevent the contaminations. | RTT-1: 22ul/aspiration (15) RTT-2: 22ul/aspiration (15) Current Contam. Settings use Clean 2 only |  |  | negligible, only if Fe before Mg |  |  |  |  | 525.6
| Reagent Probe Wash1 | Alkaline Detergent for preventing contamination of RPP and RRV (Straight) General Use | RTT-1, 2 (No.42) | two times each probe start/end of each run: 120 uL/ run | uL per run if discontinuous |  |  |  | 120 |  |  | 525.6
2 | Reagent Probe Wash2 | Acid Detergent for preventing contamination of RPP and RRV (Straight) Contamination Settings | RTT-1, 2 (No.43) | RTT-1, 2: [Clean 2 settings] * 88ul for each aspiration (60) * This aspirated detergent will be dispensed to cuvette with 132ul (90)of pure water when System makes it to prevent the contaminations. | RTT-1: 88ul/aspiration (60) RTT-2: 88ul/aspiration(60) Maximun: 120uL each MG test |  |  | 120 |  |  |  | #REF! | #REF!
| Reagent Probe Wash2 | Acid Detergent for preventing contamination of RPP and RRV (Straight) General Use | RTT-1, 2 (No.43) | two times each probe start/end of each run: 240 uL/ run | uL per run if discontinuous |  |  |  | 240 |  |  | 1051.2
3 | Reagent Probe Wash3 | Detergent for Weekly Wash (5% Dilution) | CTT | Ancillary reagent usage was estimated for Marketing | CTT : 300ul/WASH3 as undiluted 1.25*231=288.75 or 500uL | [5 empty] | 500 | 26 | 74.048
| Reagent Probe Wash3 | Detergent for Weekly Wash (5% Dilution) | RTT-1, 2 | RTT-1, 2: 231 cuvettes * 2.0ul(=40ul*5%) for each aspiration * 40ul of 5% diluted solution are aspirated. * 480ul(=2.0ul*240cuvettes) for eash WASH2 | as undiluted RTT1 480uL/Wash3 RTT2 480uL/Wash3 2*2*231=924 | [5 empty] | 924 | 48.048
4 | Cuvette Wash | Alkaline Detergent for RRV Cuvette Wash on WUD (Straight) General Use | 2L Bottle | * 76.6ul to 90.0ul for each cycle * General Analysis: 4.5sec./cycle 76.6ul*800T/H = 61.28ml/H 90.0ul*800T/H = 72.00ml/H 61.28ml to 72.00ml for 1 hour Analysis | 76.6 - 90.0ul/cycle use 85uL as average *General Analysis: 61.28 - 72.00 ml/H | 85 |  |  | 3400 |  |  | #REF! | #REF!
| Cuvette Wash | Daily WASH1 (1 cycle: H2O) |  | * WASH1 (1cycle): 4.5sec./cycle 76.6ul*39min*60sec./4.5sec. = 39.832ml 90.0ul*39min*60sec./4.5sec. = 46.800ml | *WASH1 (1cycle): 39.832 - 46.800 ml/WASH1 85uL * 231 cuvettes = 19635 plus cln for assumed emerg stop |  |  |  |  | 44200 |  | 16133
| Cuvette Wash | Daily WASH2 (2 cycles: H2O, 10%CW)) |  | * WASH2 (2cycle): 4.5sec./cycle 76.6ul*57min*60sec./4.5sec. = 58.216ml 90.0ul*57min*60sec./4.5sec. = 68.400ml | *WASH2 (2cycle): 58.216 - 68.400 ml/WASH2 2* 85uL * 231 cuvettes = 39270 plus cln for assumed emerg stop |  |  |  |  | 64600 |  | 23579
| Cuvette Wash | Weekly Cell Blank (1 cycle) |  | * Cell Blank: 3.0sec./cycle 76.6ul*17min*60sec./3.0sec. = 26.044ml 90.0ul*17min*60sec./3.0sec. = 30.600ml | *Cell Blank: 26.044 - 30.600 ml/CB 85uL * 231 cuvettes = 19635 plus cln for assumed emerg stop | [5 empty] | 28900 | 1502.8
| Cuvette Wash | Detergent for daily Wash (10% Dilution) | CTT | CTT: * 2.5ul(=25ul*10%) for each aspiration * 25ul of 5% diluted solution are aspirated. * 600ul(=2.5ul*240cuvettes) for each WASH2 | as undiluted CTT : 600ul/WASH2 2.5*231 = 577.5 |  |  |  |  | 577.5 |  | 210.7875
| Cuvette Wash | Detergent for daily Wash (10% Dilution) | RTT-1, 2 (No.44) | RTT-1, 2: * 4.0ul(=40ul*10%) for each aspiration * 40ul of 10% diluted solution are aspirated and will be dispensed to cuvette with 40ul of pure water. * 960ul(=4.0ul*240cuvettes) for each WASH2 | as undiluted RTT-1: 960ul/WASH2 RTT-2: 960ul/WASH2 2*4*231=1848 |  |  |  |  | 1848 |  | 674.52
5 | Cuvette Conditioner | Conditioner for RRV Cuvette Wash on WUD (Straight) | 2L Bottle | * 30.0ul to 46.7ul for each cycle use 40uL as average * General Analysis: 4.5sec./cycle 30.0ul*800T/H = 24.00ml/H 46.7ul*800T/H = 38.08ml/H 61.28ml to 72.00ml for 1 hour Analysis | 30.0 - 46.7ul/cycle *General Analysis: 24.0 - 38.08 ml/H use 40uL as average 40/ test | 40 |  |  | 1600 |  |  | #REF! | #REF!
| Cuvette Conditioner | Daily WASH1 |  | * WASH1 (1cycle): 4.5sec./cycle 30.0ul*39min*60sec./4.5sec. = 15.600ml 46.7ul*39min*60sec./4.5sec. = 24.284ml | *WASH1 (1cycle): 15.600 - 24.284 ml/WASH1 use 40uL as average 40*231=9240 |  |  |  |  | 20800 |  | 7592
| Cuvette Conditioner | Daily WASH2 |  | * WASH2 (2cycle): 4.5sec./cycle 30.0ul*57min*60sec./4.5sec. = 22.800ml 46.7ul*57min*60sec./4.5sec. = 35.492ml | *WASH2 (2cycle): 22.800 - 35.492 ml/WASH2 use 40uL as average 40*231=9240 |  |  |  |  | 30400 |  | 11096
| Cuvette Conditioner | Weekly Cell Blank |  | * Cell Blank: 3.0sec./cycle 30.0ul*17min*60sec./3.0sec. = 10.200ml 46.7ul*17min*60sec./3.0sec. = 15.878ml | *Cell Blank: 10.200 - 15.878 ml/CB use 40uL as average 40*231=9240 | [5 empty] | 9240 | 480.48
6 | Incubation Bath Oil | Circulating Oil for Reaction Bath (Straight) | 2L Bottle | * Baseline: under 100ml for each day * Average : 40ml for each day | Baseline: under 100ml/day Average : 40ml/day |  |  |  |  | 40000 |  | 14600 | 14600
7 | Lamp Coolant | Circulating Anticorrosion Liquid for Cooling of Detector Lamp (5% Dilution) | Lamp-cooling Liquid Tank | * This depends on environment of installed place. Small usage rate, don't calculate? | --- | [7 empty] | 0
8 | ISE Buffer | Buffer of ISE (Straight) Continuous | 2L Bottle | In case of Continuous Analysis (Waiting time: 0 sec. to less than 5 sec.): * 3.00ml for each Serum sample * 4.50ml for each Urine Sample assume no batch ISE's, and >2 chems with each ISE sample | Continuous Analysis (0 sec. - 5 sec.): * 3.00ml/Serum Sample * 4.50ml/Urine Sample |  | 0 | [5 empty] | #REF!
| ISE Buffer | Discontinuous Serum |  | In case of Discontinuous Analysis: It depends on waiting time from last analysis. * Waiting time: 5 sec. to less than 3 min. Prewash 2times(1.50ml) + Analysis: Serum: 4.50ml Urine: 6.00ml * Waiting time: 3 min. to less than 10 min. Prewash 3times(2.25ml) + Analysis Serum: 5.25ml Urine: 6.75ml * Waiting time: 10 min. and over Prewash 4times(3.00ml) + Analysis Serum: 6.00ml Urine: 7.50ml | Discontinuous Serum Analysis: * Waiting time: 5 sec. - 3 min.: 4.50ml 3 min. - 10 min.: 5.25ml 10 min. - : 6.00ml use 5 sec to 3min delay as average |  | 4500 |  |  |  |  | #REF!
| ISE Buffer | Discontinuous Urine |  | In case of Discontinuous Analysis: It depends on waiting time from last analysis. * Waiting time: 5 sec. to less than 3 min. Prewash 2times(1.50ml) + Analysis: Serum: 4.50ml Urine: 6.00ml * Waiting time: 3 min. to less than 10 min. Prewash 3times(2.25ml) + Analysis Serum: 5.25ml Urine: 6.75ml * Waiting time: 10 min. and over Prewash 4times(3.00ml) + Analysis Serum: 6.00ml Urine: 7.50ml | Discontinuous Urine Analysis: * Waiting time: 5 sec. - 3 min.: 6.00ml 3 min. - 10 min.: 6.75ml 10 min. - : 7.50ml use 5 sec to 3min delay as average |  | 6000 |  |  |  |  | 0
| ISE Buffer | Prime |  | 0.75mL per hour ? | 0.75ml * idle time/day= daily assume 12 hr idle per day |  |  |  |  | 6000 |  | 2190
| ISE Buffer | Wash Electrode |  | In case of Wash Electrode and WASH: * 4.50ml for each Wash Electrode done daily (with Wash 2 if configured) | Wash Electrode and WASH: * 4.50ml/Wash Electrode |  |  |  |  | 4500 |  | 1642.5
| ISE Buffer | Line Wash |  | In case of ISE line wash: * 7.50ml for each ISE line wash | ISE line wash: * 7.50ml/ISE line wash consider negligible | [6 empty] | 30
9 | ISE Detergent Solution | Detergent for Wash Electrode and WASH (Straight) | CTT | * 26ul for each aspiration * 52ul(=26ul*2 times) for each Wash Electrode | 52ul/Wash Electrode and WASH do daily (assume 100uL/day or top-off cup) |  |  |  |  | 500 |  | 182.5 | 182.5
10 | ISE Serum Standard H | Standard for ISE Calibration (Neat) | CTT | * 26ul for each aspiration * One Calibration makes 3 times to 8 times aspirations. 3 times: 78ul (=26*3) 8 times: 208ul (=26*8) | 78 - 208ul/Calibration (pour-out is not reused, so each cal will use 1 cup or ~1mL) |  |  |  |  | 500 |  | 365 | 365
11 | ISE Serum Standard L | Standard for ISE Calibration (Neat) | CTT | * 26ul for each aspiration * One Calibration makes 3 times to 8 times aspirations. 3 times: 78ul (=26*3) 8 times: 208ul (=26*8) | 78 - 208ul/Calibration (pour-out is not reused, so each cal will use 1 cup or ~1mL) |  |  |  |  | 500 |  | 365 | 365
12 | ISE Urine Standard H | Standard for ISE Calibration (Neat) | CTT | * 26ul for each aspiration * One Calibration makes 3 times to 8 times aspirations. 3 times: 78ul (=26*3) 8 times: 208ul (=26*8) | 78 - 208ul/Calibration (pour-out is not reused, so each cal will use 1 cup or ~1mL) |  |  |  |  | 500 |  | 365 | 365
13 | ISE Urine Standard L | Standard for ISE Calibration (Neat) | CTT | * 26ul for each aspiration * One Calibration makes 3 times to 8 times aspirations. 3 times: 78ul (=26*3) 8 times: 208ul (=26*8) | 78 - 208ul/Calibration (pour-out is not reused, so each cal will use 1 cup or ~1mL) |  |  |  |  | 500 |  | 365 | 365
14 | ISE Check Solution | Selectivity-check solution for Na, K (Neat) | CTT | --- not usually done by customer, just install, FSE? Don't estimate | ---
Note on K20: Bayer Corporation: DD data, new cuvettes

## URH-WORKCELL A1650 (hidden)
Customer Name
Number of Analyzers by Model
ADVIA 1650 | 1
ADVIA 2400 | 0
1
Customer Parameters | Value | Ancillary | mL/ Year | Bottles / Year |  | Packaging | SPC/ Package | Cost/year
Tests per Year (patients, incl ISE) | #REF! | Probe Wash 1 | 254.688 | 2 | 1 | 5x250mL | 25.88 | 25.88 | 2 | 3 | 4 | 1
% dilns & reruns | 0 | Probe Wash 2 | 614.112 | 3 | 1 | 5x250mL | 35.33 | 35.33 | 2 | 3 | 4 | 1
Total Chem Tests/ year | #REF! | Probe Wash 3 | 158.6 | 1 | 1 | 1x500mL | 16.73 | 16.73 |  |  |  | 1
Mg Tests per year | #REF! | Cuvette Wash | #REF! | #REF! | #REF! | 1x2L | 13.22 | #REF! |  |  |  | #REF!
-or- %Mg Tests | #REF! | Cuvette Conditioner | #REF! | #REF! | #REF! | 1x1L | 19.24 | #REF! |  |  |  | #REF!
ISE tests per year | #REF! | ISE Buffer | #REF! | #REF! | #REF! | 1x2L | 28.42 | #REF! |  |  |  | #REF!
serum ISE panel's per year | #REF! | ISE Reference Soln. | #REF! | #REF! | #REF! | 1x1L | 40.02 | #REF! |  |  |  | #REF!
urine ISE panel's per year | 0 | Incubation Bath Oil | 18250 | 7 | 7 | 1x2.9L | 279 | 1953 |  |  |  | 7
1
Shifts per Day | 2 | Lamp Coolant |  | 1 | 1 | ~50mL | 23.1 | 23.1 |  |  |  | 1
Runs per Analyzer/Day | 12 |  |  |  | 1
Idle Hours per Analyzer/Day | 8 | ISE Detergent | 31.2 | 1 | 1 | 2x100mL | 24.15 | 24.15 | 2 | 3 | 4 | 1
Amount of days per year | 312
TOTAL COST | [7 empty] | #REF!
Calculated cells (grey)
Enter usage data (yellow)
Bottles/year calculation includes dead volume
ADDITIONAL CONSUMABLES (not dependent on test volume)
|  |  | Recommended Interval | # Used/ Year
Cuvettes |  |  | 63 | 63 |  | 1 segment | 82 | 5166 |  |  |  | 20
Lamp |  |  | 4 | 4 |  | each | 77 | 308 |  |  |  | 4
Electrodes (assume change every 3 months) |  |  | 12 | 12 |  |  | 290 | 3480 |  |  |  | 4
Ref Electrode |  |  | 1 | 1 |  |  | 676.36 | 676.36 |  |  |  | 1
TOTAL | [7 empty] | #REF!

## STAND ALONE ADVIA 1650 (hidden)
Customer Name
Number of Analyzers by Model
ADVIA 1650 | 1
ADVIA 2400
1
Customer Parameters | Value | Ancillary | mL/ Year | Bottles / Year |  | Packaging | SPC/ Package | Cost/year
Tests per Year (patients, incl ISE) | 700000 | Probe Wash 1 | 401.64 | 2 | 1 | 5x250mL | 25.88 | 25.88 | 2 | 3 | 4
% dilns & reruns | 0.05 | Probe Wash 2 | 1032.36 | 5 | 2 | 5x250mL | 35.33 | 70.66 | 2 | 3 | 4
Total Chem Tests/ year | -167100 | Probe Wash 3 | 158.6 | 1 | 1 | 1x500mL | 16.73 | 16.73
Mg Tests per year | 7000 | Cuvette Wash | 19625.243 | 10 | 10 | 1x2L | 13.22 | 132.2
-or- %Mg Tests | 0.01 | Cuvette Conditioner | 11907.9 | 12 | 12 | 1x1L | 19.24 | 230.88
ISE tests per year | 902100 | ISE Buffer | 3437472.213 | 1719 | 1719 | 1x2L | 28.42 | 48853.98
serum ISE panel's per year | 300000 | ISE Reference Soln. | 340489.8187 | 341 | 341 | 1x1L | 40.02 | 13646.82
urine ISE panel's per year | 700 | Incubation Bath Oil | 18250 | 7 | 7 | 1x2.9L | 279 | 1953
1
Shifts per Day | 3 | Lamp Coolant | 1 | 1 | 1 | ~50mL | 23.1 | 23.1
Runs per Analyzer/Day | 18 |  |  |  | 1
Idle Hours per Analyzer/Day | 1.87 | ISE Detergent | 36.5 | 1 | 1 | 2x100mL | 24.15 | 24.15 | 2 | 3 | 4
Amount of days per year | 365
TOTAL COST | [7 empty] | 64977.4
Calculated cells (grey)
Enter usage data (yellow)
Bottles/year calculation includes dead volume
ADDITIONAL CONSUMABLES (not dependent on test volume)
|  |  | Recommended Interval | # Used/ Year
Cuvettes |  |  | 63 | 63 |  | 1 segment | 82 | 5166
Lamp |  |  | 4 | 4 |  | each | 77 | 308
Electrodes (assume change every 3 months) |  |  | 12 | 12 |  |  | 290 | 3480
Ref Electrode |  |  | 1 | 1 |  |  | 676.36 | 676.36
TOTAL | [7 empty] | 74607.76

## Stand Alone Calculations (hidden)
Probe Wash 1
Run type | RTT 1 (uL) | RTT 2 (uL) | Yearly Amount (mL)
Variable amount per run/day | 26 | 26 | 341.64
Dead volume/year (5,000 mL/2months) | 30000 | 30000 | 60
Total/year |  |  | 401.64
Probe Wash 2
Run type | RTT 1 (uL) | RTT 2 (uL) | Yearly Amount (mL)
Variable amount per run/day | 74 | 74 | 972.36
Dead volume/year (5,000 mL/2months) | 30000 | 30000 | 60
Total/year |  |  | 1032.36
Probe Wash 3
Run type | RTT 1 (mL) | RTT 2 (mL) | CTT (mL) | Yearly Amount (mL)
WASH 2 | 23 | 23 | 7.2
Amount used in 52 weeks | 1196 | 1196 | 374.4 | 2766.4
Dead volume/year (3,120 mL/week) | 156 | 156 | 93.6 | 405.6
Total/year | 1352 | 1352 | 468 | 3172
Total Undiluted | 67.6 | 67.6 | 23.4 | 158.6
Cuvette Wash
Run type | RTT 1 (mL) | RTT 2 (mL) | CTT (mL) | Miscellaneous (mL)
WASH 2 (Shutdown) | 23 | 23 | 7.2
Amount used in 313 days | 7199 | 7199 | 2253
Dead volume in 1 year | 939 | 939
10-mL at CTT position # 49 waste |  |  |  | 563
Weekly Cell Blank/year |  |  |  | 962
Wash 2 for RRV |  |  |  | 14555
Wash 3 for RRV |  |  |  | 12363
Waste Amount |  |  |  | 3000
Sub Total | 8161 | 8161 | 2260.2 | 18613.643
Amount used before 1st sample |  |  |  | 11037.6
Variable Amount per Volume of Test |  |  |  | -10026
Sub Total |  |  |  | 1011.6
Total Undiluted |  |  |  | 19625.243
Cuvette Conditioner
Run type | Amount used (mL) | Yearly amount (mL)
Wash 2 | 12.5 | 4562.5
Wash 3 | 10.5 | 3832.5
Weekly Cell Blank | 5 | 260
Amount Discarded | 250 | 3000
Sub Total |  | 11655
Amount used before 1st sample |  | 2759.4
Variable Amount per Volume of Test |  | -2506.5
Sub Total |  | 252.9
Total Undiluted |  | 11907.9
ISE Solutions
Run type | Amount used (mL) | Yearly amount (mL)
ISE Detergent During Wash 2 | 0.1 | 36.5
Incubation Bath Oil | 50 | 18250
ISE Reference Solution
Start up Prime | 0.4 | 146
Dummy Cycle | 0.3 | 1971
Sub Total |  | 2117
ISE Reference Solution (auto-prime) | 0.1 | 68.255
Variable ISE Solution per sample | 0.3 | 270630
Sub Total |  | 270698.255
ISE Reference Solution Variable waste | 250 | 67674.5638
Total ISE Reference Solution |  | 340489.8187
ISE Buffer Solution
ISE Buffer Start Up Prime | 3.6 | 1314
ISE Buffer Solution (auto-prime) | 0.9 | 614.295
Dummy Cycle | 2.7 | 17739
Sub Total |  | 19667.295
Variable ISE Buffer Solution per sample | 2.7 | 2435670
Sub Total |  | 2455337.295
ISE Buffer Solution Variable waste | 400 | 982134.918
Total ISE Buffer Solution |  | 3437472.213

## URH-Workcell Calc (hidden)
Probe Wash 1
Run type | RTT 1 (uL) | RTT 2 (uL) | Yearly Amount (mL)
Variable amount per run/day | 26 | 26 | 194.688
Dead volume/year (5,000 mL/2months) | 30000 | 30000 | 60
Total/year |  |  | 254.688
Probe Wash 2
Run type | RTT 1 (uL) | RTT 2 (uL) | Yearly Amount (mL)
Variable amount per run/day | 74 | 74 | 554.112
Dead volume/year (5,000 mL/2months) | 30000 | 30000 | 60
Total/year |  |  | 614.112
Probe Wash 3
Run type | RTT 1 (mL) | RTT 2 (mL) | CTT (mL) | Yearly Amount (mL)
WASH 2 | 23 | 23 | 7.2
Amount used in 52 weeks | 1196 | 1196 | 374.4 | 2766.4
Dead volume/year (3,120 mL/week) | 156 | 156 | 93.6 | 405.6
Total/year | 1352 | 1352 | 468 | 3172
Total Undiluted | 67.6 | 67.6 | 23.4 | 158.6
Cuvette Wash
Run type | RTT 1 (mL) | RTT 2 (mL) | CTT (mL) | Miscellaneous (mL)
WASH 2 (Shutdown) | 23 | 23 | 7.2
Amount used in 313 days | 7199 | 7199 | 2253
Dead volume in 1 year | 939 | 939
10-mL at CTT position # 49 waste |  |  |  | 563
Weekly Cell Blank/year |  |  |  | 962
Wash 2 for RRV |  |  |  | 14555
Wash 3 for RRV |  |  |  | 12363
Waste Amount |  |  |  | 3000
Sub Total | 8161 | 8161 | 2260.2 | 50025.2
Amount used before 1st sample |  |  |  | 6289.92
Variable Amount per Volume of Test |  |  |  | #REF!
Sub Total |  |  |  | #REF!
Total Undiluted |  |  |  | #REF!
Rate per Hour Usage |  |  |  | 0
Cuvette Conditioner
Run type | Amount used (mL) | Yearly amount (mL)
Wash 2 | 12.5 | 3900
Wash 3 | 10.5 | 3276
Weekly Cell Blank | 5 | 260
Amount Discarded | 250 | 3000
Sub Total |  | 10436
Amount used before 1st sample |  | 1572.48
Variable Amount per Volume of Test |  | #REF!
Sub Total |  | #REF!
Total Undiluted |  | #REF!
Rate per Hour Usage |  | 0
ISE Solutions
Run type | Amount used (mL) | Yearly amount (mL)
ISE Detergent During Wash 2 | 0.1 | 31.2
Incubation Bath Oil | 50 | 18250
ISE Reference Solution
Start up Prime | 0.4 | 124.8
Dummy Cycle | 0.3 | -1123.2
ISE Reference Solution (auto-prime) | 0.1 | 249.6
Sub Total |  | -748.8
Variable ISE Solution per sample | 0.3 | #REF!
Sub Total |  | #REF!
ISE Reference Solution Variable waste | 250 | #REF!
Total ISE Reference Solution |  | #REF!
ISE Buffer Solution
ISE Buffer Start Up Prime | 3.6 | 1123.2
ISE Buffer Solution (auto-prime) | 0.9 | 2246.4
Dummy Cycle | 2.7 | -10108.8
Sub Total |  | -6739.2
Variable ISE Buffer Solution per sample | 2.7 | #REF!
Sub Total |  | #REF!
ISE Buffer Solution Variable waste | 400 | #REF!
Total ISE Buffer Solution |  | #REF!

## Advia Centaur
Importación Regular
ADVIA CENTAUR XP / CP
10309958 | TSH 100 tests Advia Centaur XP / CP | NO
10309959 | TSH 500 tests Advia Centaur XP / CP | SI
10282378 | TSH-3 ultra 100 tests Advia Centaur XP / CP | NO
10282379 | TSH-3 ultra 500 tests Advia Centaur XP / CP | SI
10309960 | T4 100 tests Advia Centaur XP / CP | SI
10309961 | T4 500 tests Advia Centaur XP / CP | SI
10285733 | T3 x 80 tests Advia Centaur XP / CP | SI
10285732 | T3 x 400 tests Advia Centaur XP / CP | SI
10282218 | FT4 x 50 tests Advia Centaur XP / CP | SI
10282219 | FT4 x 250 tests Advia Centaur XP / CP | SI
10310317 | FT3 50 tests Advia Centaur XP / CP | SI
10310318 | FT3 250 tests Advia Centaur XP / CP | NO
10630886 | Anti-TPO 100 tests Advia Centaur XP / CP | SI
10630887 | Anti-TPO 500 tests Advia Centaur XP / CP | SI
10492398 | Anti-TG 100 tests Advia Centaur XP / CP | SI
10492399 | Anti-TG 500 tests Advia Centaur XP / CP | NO
10309966 | TUp 50 tests Advia Centaur XP / CP | NO
10309967 | TUp 250 tests Advia Centaur XP / CP | NO
10309971 | LH 60 tests Advia Centaur XP / CP | SI
10309972 | LH 300 tests Advia Centaur XP / CP | SI
10309973 | FSH 100 tests Advia Centaur XP / CP | SI
10309974 | FSH 500 tests Advia Centaur XP / CP | SI
10309975 | Prolactina 50 tests Advia Centaur XP / CP | SI
10309976 | Prolactina 250 tests Advia Centaur XP / CP | SI
10490889 | eEstradiol x 100 Advia Centaur XP / CP | SI
10491445 | eEstradiol x 500 Advia Centaur XP / CP | SI
10310305 | Progesterona 50 tests Advia Centaur XP / CP | SI
10315522 | Progesterona 250 tests Advia Centaur XP / CP | SI
10309989 | Testosterona 50 tests Advia Centaur XP / CP | SI
10322623 | Testosterona 250 tests Advia Centaur XP / CP | SI
10308984 | HCG 50 tests Advia Centaur XP / CP | SI
10308985 | HCG 250 tests Advia Centaur XP / CP | SI
10310292 | PSA e 100 tests Advia Centaur XP / CP | SI
10310293 | PSA e 500 tests Advia Centaur XP / CP | SI
10310381 | PSA Complejado 100 tests Advia Centaur XP / CP | NO
10283414 | Free PSA 50 tests Advia Centaur XP | SI
10491379 | CA 19-9 50 tests Advia Centaur XP / CP | SI
10491244 | CA 19-9 250 tests Advia Centaur XP / CP | SI
10327620 | CA 15-3 100 tests Advia Centaur XP / CP | SI
10312831 | CA 15-3 500 tests Advia Centaur XP / CP | NO
10310443 | CA 125 II 100 tests Advia Centaur XP / CP | SI
10315696 | CA 125 II 500 tests Advia Centaur XP / CP | NO
10308994 | Her-2 Neu 50 tests Advia Centaur XP / CP | NO
10309977 | CEA 100 tests Advia Centaur XP / CP | SI
10309978 | CEA 500 tests Advia Centaur XP / CP | SI
10309979 | AFP 100 tests Advia Centaur XP / CP | SI
10309980 | AFP 500 tests Advia Centaur XP / CP | NO
10309970 | VB12 100 tests Advia Centaur XP / CP | SI
10330218 | VB12 500 tests Advia Centaur XP / CP | NO
10310308 | Folato 100 tests Advia Centaur XP / CP | SI
10325366 | Folato 500 tests Advia Centaur XP / CP | NO
10309968 | Ferritina 50 tests Advia Centaur XP / CP | SI
10309969 | Ferritina 250 tests Advia Centaur XP / CP | SI
10309044 | BNP 100 tests Advia Centaur XP / CP | SI
10309982 | CK MB 100 tests Advia Centaur XP / CP | NO
10326496 | CK MB 500 tests Advia Centaur XP / CP | NO
10317708 | TnI Ultra 100 tests Advia Centaur XP / CP | SI
10317709 | TnI Ultra 500 tests Advia Centaur XP / CP | NO
10310277 | Mioglobina 50 tests Advia Centaur XP / CP | SI
10329242 | Mioglobina 250 tests Advia Centaur XP / CP | NO
10310374 | Homocisteína 100 tests Advia Centaur XP / CP | SI
10310375 | Homocisteína 500 tests Advia Centaur XP / CP | NO
10310439 | Insulina 100 tests Advia Centaur XP / CP | SI
10310448 | Peptido C 100 tests Advia Centaur XP / CP | SI
10492368 | iPTH 100 tests Advia Centaur XP / CP | SI
10492369 | iPTH 500 tests Advia Centaur XP / CP | SI
10282227 | DHEA SO4 50 tests Advia Centaur XP / CP | SI
10699201 | Vitamin D 100 tests Estandarizada Advia Centaur XP | SI
10699533 | Vitamin D 500 tests Estandarizada Advia Centaur XP | NO
10309985 | IgE 50 tests Advia Centaur XP / CP | SI
10309986 | IgE 250 tests Advia Centaur XP / CP | SI
10994924 | Enhanced Cortisol 50 tests Advia Centaur XP / CP | SI
10994926 | Enhanced Cortisol 250 tests Advia Centaur XP / CP | SI
10283087 | SHBG 50 tests Advia Centaur XP / CP | SI
10322033 | Carbamazepina 50 tests Advia Centaur XP / CP | SI
10322035 | Carbamazepina 250 tests Advia Centaur XP / CP | NO
10322090 | Fenobarbital 50 tests Advia Centaur XP / CP | SI
10310328 | Fenitoína 50 tests Advia Centaur XP / CP | SI
10328805 | Fenitoína 250 tests Advia Centaur XP / CP | NO
10309981 | Digoxina 50 tests Advia Centaur XP / CP | SI
10322129 | Digoxina 250 tests Advia Centaur XP / CP | NO
10310301 | Digitoxina 50 tests Advia Centaur XP / CP | NO
10329308 | Digitoxina 250 tests Advia Centaur XP / CP | NO
10322154 | Gentamicina 50 tests Advia Centaur XP / CP | NO
10310434 | Teofilina 100 tests Advia Centaur XP / CP | SI
10328931 | Teofilina 500 tests Advia Centaur XP / CP | NO
10322215 | Acido Valproico 50 tests Advia Centaur XP / CP | SI
10322218 | Acido Valproico 250 tests Advia Centaur XP / CP | NO
10310300 | Vancomicina 50 tests Advia Centaur XP / CP | SI
10378883 | Procalcitonina 100 tests Advia Centaur XP / CP | NO
10309022 | EHIV 200 tests Advia Centaur XP / CP | SI
10283020 | cHIV(XUS) Advia Centaur XP / CP 100t | SI
10309061 | HCV 200 tests Advia Centaur XP / CP | SI
10309057 | HBsAg 200 tests Advia Centaur CP | SI
10492138 | HBsAgII 200 tests Advia Centaur XP | SI
10309058 | HBsAg Confirmatorio Advia Centaur XP / CP | NO
10286268 | Anti-HBs2 200 tests Advia Centaur XP / CP | SI
10309508 | Anti-HBc 200 tests Advia Centaur XP / CP | SI
10308978 | IgM anti-HBc 100 tests Advia Centaur XP / CP | SI
10334094 | aHBe 50 tests Advia Centaur XP | SI
10291295 | HBeAg 50 tests Advia Centaur XP | SI
10322676 | Anti-HVA 100 tests Advia Centaur XP / CP | SI
10309083 | IgM anti-HVA 100 tests Advia Centaur XP / CP | SI
10492493 | Sifilis Advia Centaur XP / CP 200T | NO
10310283 | Rubeola IgG 100 tests Advia Centaur XP / CP | SI
10310285 | Rubeola IgM 50 tests Advia Centaur XP / CP | SI
10310321 | Toxoplasma IgG 100 tests Advia Centaur XP / CP | SI
10310320 | Toxoplasma IgM 50 tests Advia Centaur XP | SI
CALIBRADORES ADVIA CENTAUR XP / CP
10285903 | Calibrador A (2x2x5ml) | SI
10308986 | Calibrador B (2x2x5ml) | SI
10311568 | Calibrador C (2x2x5ml) | SI
10311569 | Calibrador D (2x2x2ml) | SI
10309079 | Calibrador E Enhanced (2x2x2ml) | SI
10322159 | Calibrador J (Genta) (2x2x2ml) | NO
10630932 | Calibrador O (2x2x1ml) | SI
10310295 | Calibrador Q (2x2x2ml) | SI
10309996 | Calibrador U (2x2x2ml) | NO
10311377 | Calibrador V (2x2x2ml) | NO
10310383 | Calibrador Y (2x2x2ml) | NO
10322039 | Calibrador Z (2x2x5ml) | SI
10361919 | Calibrador Free PSA | SI
10311570 | Calibrador CK MB (2x2x2ml) | NO
10311391 | Calibrador Fenitoína (2x2x2ml) | SI
10322092 | Calibrador Fenobarbital (2x2x2ml) | SI
10310376 | Calibrador HCY (2x2x2ml) | SI
10310438 | Calibrador Insulina (2x2x2ml) | SI
10310445 | Calibrador Peptido C (2x2x2ml) | SI
10630915 | Calibrador 1 (2x2x1ml) | SI
10379810 | Calibrador 30 para Estradiol Nueva Version (2x2x2ml) | SI
10322222 | Calibrador 28 (2x2x2ml) | SI
10492394 | Calibrador 56 (2x2x2ml) | SI
10492491 | Calibrador Ige 80 | SI
10309047 | Calibrador BNP (2x2x2ml) | SI
10310432 | Calibrador 27 Teofilina (2x2x2ml) | SI
10310444 | Calibrador CA 125 II (2x2x2ml) | SI
10329796 | Calibrador CA 15-3 (2x2x2ml) | SI
10308993 | Calibrador Her-2 Neu (2x2x2ml) | NO
10285318 | Calibrador SHBG | SI
CONTROLES ADVIA CENTAUR XP / CP
10282764 | Bio Rad Lipochek Immunoassay Plus (Trilevel 3 x 4 x 5ml) | SI
10710159 | Biorad Liquicheck Tumor Marker Control - Trilevel MiniPack (3 x 2ml) | SI
10282768 | Biorad Liquicheck Cardiac Markers Plus LT - Nivel 1 (6 x 3 ml) | SI
10282769 | Biorad Liquicheck Cardiac Markers Plus LT - Nivel 2 (6 x 3 ml) | SI
10282770 | Biorad Liquicheck Cardiac Markers Plus LT - Nivel 3 (6 x 3 ml) | SI
10308992 | Control Her-2 Neu (2x2x2ml) | NO
10630936 | Control anti-TPO 1,2 | SI
10310387 | Control Calidad cPSA | NO
10630917 | Control para Anti-TG 1,2 | SI
10309046 | Control para BNP | SI
10492378 | Control iPTH | SI
10378884 | Control para Procalcitonina | NO
10310284 | Control Rubeola IgG | SI
10310286 | Control Rubeola IgM | SI
10310325 | Control Toxoplasma IgG | SI
10310326 | Control Toxoplasma IgM | SI
10309010 | Control HIV Enhanced 3x2x7ml | SI
10283022 | CENTAUR CHIV QC KIT | SI
10309062 | Control HCV 2x2x7ml | SI
10309059 | Control HBsAg 2x2x10ml | SI
10283088 | Control Anti-HBs2 2x2x7ml | SI
10308979 | Control IgM anti-HBc 2x2x7ml | SI
10309509 | Control Anti-HBc 2x2x7ml | SI
10334181 | Control aHbe 2 x 10 ml | SI
10291296 | Control HBeAg 2 x 10 ml | SI
10309084 | Control IgM anti-HAV 2x2x7ml | SI
10322682 | Control Anti-HAV 2x2x7ml | SI
10492616 | Sifilis QC | NO
10699200 | Control Vitamina D ESTANDARIZADA | SI
CONSUMIBLES ADVIA CENTAUR XP
10310026 | Centaur:Reactivo 1&2 x 5000 | SI
CONSUMIBLES ADVIA CENTAUR CP
10313526 | Centaur CP: Reactivo 1 & 2 | SI
10316997 | Centaur CP APW2 (Testosterona / estradiol) | SI
CONSUMIBLES COMUNES ADVIA CENTAUR XP/ CP
10309547 | Probe Tips x 6480 | SI
10309546 | Cuvettes (3000/box) | SI
10309545 | Sample Cups (1500/box) | SI
10311563 | Sample Cup Caps (1000/pack) | SI
10309954 | T3/T4/VB12 Releasing Agent 6 Pack | SI
10310330 | V12 DTT/Releasing Agent | SI
10310382 | DTT / Agente liberador Folato | SI
10310272 | Rvo. Pretratamiento BR | NO
10310384 | Rvo. Pretratamiento PSA complejado | NO
10310032 | Wash 1 (Lavado 1) 2x1500 ml | SI
10309060 | APW1 | SI
10316877 | Reactivo Auxiliar Fólico Glóbulos Rojos (4x30ml) | SI
10309942 | MultiDiluyente 1 | SI
10309943 | MultiDiluyente 2 | SI
10491878 | Diluyente para Estradiol | SI
10491972 | Diluyente CA 19-9 | SI
10323214 | Diluyente HCG | SI
10309952 | IgE Diluyente | SI
10309950 | Diluyente para CEA | SI
10334314 | Centaur APW 3 | SI
10310430 | Diluyente aTG | NO
10310391 | Diluyente aTPO | NO
10310358 | Solución PW (Probe Wash) | NO

## IMMULITE 1000
Importación Regular
IMMULITE 1 / IMMULITE 1000
10381630 | T3 100 tests | SI
10381626 | Free T3 100 tests | SI
10381633 | T4 100 tests | SI
10381635 | T4 500 tests | SI
10381622 | Free T4 100 tests | SI
10381620 | 3rd Generation TSH 100 tests | SI
10381627 | 3rd Generation TSH 500 tests | SI
10381637 | Rapid TSH 100 tests | SI
10381628 | Rapid TSH 500 tests | SI
10381618 | Anti-TPO 100 tests | SI
10381641 | Anti-Thyroglobulin 100 tests | SI
10381644 | Thyroglobulin 100 tests | SI
10381643 | TBG 100 tests | NO
10381619 | T Uptake 100 tests | NO
10381161 | HCG 100 tests | SI
10381164 | Free Beta HCG 100 tests | SI
10381150 | PAPP-A (Pregnancy Associated Plasma Protein) 100 tests | SI
10381153 | FSH 100 tests | SI
10381141 | LH 100 tests | SI
10381145 | Prolactin 100 tests | SI
10381132 | Estradiol 100 tests | SI
10381128 | Progesterone 100 tests | SI
10381156 | Total Testosterone 100 tests | SI
10381388 | Cortisol 100 tests | SI
10381158 | Androstenedion 100 tests | SI
10381130 | DHEA-SO4 100 tests | SI
10381168 | Unconjugated Estriol 100 tests | SI
10381415 | ACTH 100 tests | SI
10381403 | IGF-I 100 tests | SI
10381425 | IGFBP-3 100 tests | SI
10381401 | Calcitonin 100 tests | SI
10380950 | Beta-2 Microglobulin 100 tests | SI
10381433 | HGH 100 tests - Nueva Versión | SI
10381138 | SHBG (GLAE) 100 tests | SI
10381389 | DPD - Pyrilinks-D 100 tests | SI
10381396 | Osteocalcin 100 tests | SI
10381397 | Peptido C 100 tests | SI
10381429 | Insulin 100 tests | SI
10381399 | Intact-PTH Kit (PTH Molécula Intacta)(mono/poly) 100 tests | SI
10381430 | Intact-PTH Kit TURBO 50 tests | SI
10380898 | Ferritin 100 tests | SI
10381385 | Albumin (Microalbumina en Orina) 100 tests | SI
10487627 | EPO 100 tests | SI
10380900 | Vitamin B12 100 tests | SI
10381162 | AFP 100 tests | SI
10380945 | CEA 100 tests | SI
10380960 | PSA 100 tests | SI
10380949 | PSA 500 tests | SI
10380956 | 3rd Generation PSA 100 tests | SI
10380967 | Free PSA 100 tests | SI
10380953 | PAP (Fosfatasa Acida Prostatica) 100 tests | NO
10380948 | BR-MA (CA 15.3) 100 tests | SI
10380969 | OM-MA (CA 125) 100 tests | SI
10380965 | GI-MA (19.9) 100 tests | SI
10381517 | Digoxin 100 tests | NO
10381508 | Digitoxina 100 tests | NO
10381503 | Theophylline 100 tests | NO
10381505 | Carbamazepine 100 tests | SI
10381510 | Phenytoin 100 tests | SI
10381509 | Valproic Acid 100 tests | SI
10381512 | Phenobarbital 100 tests | SI
10381507 | Nicotine Metabolite 100 tests | NO
10286287 | Proteina C Reactiva (Ultrasensible) 100 tests | SI
10381029 | Homocystein 100 tests | SI
10381016 | CK-MB 100 tests | NO
10381022 | CK-MB Turbo x 100 tests | NO
10381023 | Myoglobin 100 tests | NO
10381026 | Troponina I 100 tests | SI
10381014 | Troponina I Turbo 100Tests | NO
10381011 | NT Pro BNP Turbo 100 tests | NO
10381012 | Dimero D Turbo 100 tests | NO
10381268 | Toxoplasma Quant. IgG 100 tests | SI
10381288 | Toxoplasma IgM u-Capture 100 tests | SI
10381292 | Rubella Quant. IgG 100 tests | SI
10381282 | Rubella IgM 100 tests | SI
10381274 | Herpes 1&2 IgG Cualitativa 100 tests | NO
10381289 | Citomegalovirus IgG Cualitativa 100 tests | SI
10381265 | Citomegalovirus IgG Cualitativa 500 tests | NO
10381296 | Citomegalovirus IgM Cualitativa 100 tests | SI
10381272 | H.Pylori Semi-Quantitative IgG 100 tests | SI
10380867 | Total IgE 100 tests | SI
10380870 | ECP (Proteína Catiónica de Eosinófilos) 50 tests | NO
10380869 | IgE Especifica Inhalantes AlaTOP Allergy Screen 100 tests | NO
10381421 | Interleuquina IL-1ß 100 tests | NO
10381416 | Interleuquina IL 2R Kit (Mono/Poly) 100 tests | NO
10381411 | Interleuquina IL 6 kit (Mono/Poly) 100 tests | NO
10381407 | Interleuquina IL 8 100 tests | NO
10381417 | TNF - alfa 100 tests | NO
10381395 | LBP (Proteína Ligadora de Lipopolisacáridos) 100 tests | NO
CONTROLES IMMULITE 1 / IMMULITE 1000
10282764 | Bio Rad Lipochek Immunoassay Plus (Trilevel 3 x 4 x 5ml) | SI
10710159 | Biorad Liquicheck Tumor Marker Control - Trilevel MiniPack (3 x 2ml) | SI
10282768 | Biorad Liquicheck Cardiac Markers Plus LT - Nivel 1 (6 x 3 ml) | SI
10282769 | Biorad Liquicheck Cardiac Markers Plus LT - Nivel 2 (6 x 3 ml) | SI
10282770 | Biorad Liquicheck Cardiac Markers Plus LT - Nivel 3 (6 x 3 ml) | SI
10381639 | Thyroglobulin Recovery 2x1ml | SI
10385243 | Pyrilinks-D Control Module 2x2ml | SI
10385244 | TBG Control Module 2x2ml | SI
10385246 | Thyroglobulin Control Module 3x2ml | SI
10385316 | C-Reactive Protein Control Module 3x2ml | SI
10385317 | PAPP-A Control Module 2x1ml | SI
10385337 | LBP Control Module 2x1ml | NO
10385349 | Total IgE Control Module 2x2ml | SI
10385354 | Intact PTH Control Module 6x2ml | SI
10385358 | HCG Control Module 2x2ml | SI
10385381 | Autoantibody (TPO - TG) Control Module 2x5ml | SI
10385382 | ACTH Control Module 2x2ml | SI
10385383 | Control Calcitonin 2x3ml | SI
10385384 | EPO Control Module 3x2ml | SI
10385385 | Free Beta HCG Control Module 3x1ml | SI
10385387 | IGF-1 e IGF-BP3 Control Module 2x2ml | SI
10385388 | Albumin Control Module 2x2ml | SI
10385389 | Control Interleuquinas Module 2x5ml | NO
10385395 | 3rd Generation TSH Control Module (aprox 0,010 mUI/ml) 2x2ml | NO
10385397 | Serum Drug Control Module 2x5ml | SI
10385587 | ß-2 Microglobulin Control Module 3x1ml | SI
10385588 | PAP Control Module 3x2ml | NO
10385589 | C-Peptide Control Module 3x2ml | SI
10385339 | Control NT PRO BNP | NO
10385378 | Control Ferritina | SI
10385330 | Control para Dimero D | NO
10385331 | Control para Gastrina | SI
10385415 | Control Osteocalcina 2x2ml | SI
CONSUMIBLES Y AUXILIARES IMMULITE 1 / IMMULITE 1000
10385393 | Solución de Lavado (Probe Wash Module) 2000 tests | SI
10385394 | Sustrato (Chemiluminescent Substrate Module) 2x100 1000 tests | SI
10385183 | Tapas Copas de Muestra Immulite (Sample Cup Caps) 1000 tests | SI
10380037 | Copas de Muestra Immulite (Sample Cups) 1000 tests | SI
10385360 | Sample Cup Holder Set 1-50 | SI
10385319 | SAMPLE CUP HOLDER SET 51-100 | SI
10385348 | Dilution Cup Holder (soportes para copas de dilución) para IMM1K | SI
10385353 | Diluyente Helicobacter Pylori | SI
10387014 | Diluyente ACTH | NO
10387016 | Diluyente B2-microglobulina | SI
10387018 | Diluyente BR- MA (CA 15-3) | NO
10386977 | Diluyente HCG y BHCG | SI
10387030 | Diluyente HGH | NO
10387031 | Diluyente Albumina | SI
10386937 | Diluyente IgE | NO
10387036 | Diluyente OM-MA (CA 125) | NO
10387037 | Diluyente PAP | NO
10387039 | Diluyente Peptido C | NO
10386991 | Diluyente Progesterona | NO
10283037 | Diluyente iPTH | NO
10387040 | Diluyente Prolactina | SI
10386994 | Diluyente PSA | NO
10387041 | Diluyente SHBG | SI
10386996 | Diluyente TSH | NO
10386999 | Diluyente Vit. B12 | SI
10387047 | Diluyente CEA | SI
10387015 | Diluyente AFP | NO
10387020 | Diluyente DHEA-SO4 | NO
10283034 | Diluyente GI-MA (CA 19-9) | NO
10387026 | Diluyente FSH | NO
10387021 | Diluyente Estradiol | NO
10386981 | Diluyente Cortisol | SI
10387046 | Diluyente Tiroglobulina | SI
10386985 | Diluyente para Acido Folico | SI
10387608 | IML.1000 IGG/ IGM DIL. MODULE 100 tests | SI
10387019 | Diluyente para Calcitonina | NO
10385366 | Diluyente SHBG | SI
10387025 | Diluyente para Ferritina | NO

## IMMULITE 2000
Importación Regular
IMMULITE 2000
| 10381654 | T3 200 tests | SI
| 10381657 | T3 600 tests | SI
| 10381675 | Free T3 200 tests | SI
| 10381682 | Free T3 600 tests | NO
| 10381685 | T4 200 tests | SI
| 10381664 | T4 600 tests | SI
| 10381678 | Free T4 200 tests | SI
| 10381677 | Free T4 600 tests | SI
| 10381665 | 3rd Generation TSH 200 tests | SI
| 10381667 | 3rd Generation TSH 600 tests | SI
| 10381651 | Rapid TSH 200 tests | SI
| 10381652 | Rapid TSH 600 tests | SI
| 10381650 | Anti-TPO 200 tests | SI
| 10381649 | Anti-TPO 600 tests | SI
| 10381659 | Anti-TG 200 tests | SI
| 10381655 | Anti-TG 600 tests | SI
| 10381648 | Thyroglobulin 200 tests | SI
| 10381660 | TBG 200 tests | SI
| 10381206 | HCG 200 tests | SI
| 10381194 | HCG 600 tests | SI
| 10381175 | Free Beta HCG 200 tests | SI
| 10381213 | PAPP-A 200 tests | SI
| 10381201 | FSH 200 tests | SI
| 10381180 | FSH 600 tests | SI
| 10381211 | LH 200 tests | SI
| 10381212 | LH 600 tests | SI
| 10381200 | Prolactin 200 tests | SI
| 10381199 | Prolactin 600 tests | SI
| 10381178 | Estradiol 200 tests | SI
| 10381177 | Estradiol 600 tests | SI
| 10381181 | Progesterone 200 tests | SI
| 10381170 | Progesterone 600 tests | NO
| 10381190 | Total Testosterone 200 tests | SI
| 10381191 | Total Testosterone 600 tests | NO
| 10381476 | Cortisol 200 tests | SI
| 10381480 | Cortisol 600 tests | NO
| 10381188 | Androstenedione 200 tests | SI
| 10381193 | DHEA-SO4 200 tests | SI
| 10381192 | Unconjugated Estriol 200 tests | SI
| 10381468 | ACTH 200 tests | SI
| 10381448 | IGF-I 200 tests | SI
| 10381447 | IGFBP-3 200 tests | SI
| 10380979 | Gastrin 200 tests | SI
| 10381446 | Calcitonin 200 tests | SI
| 10381451 | HGH 200 tests - Nueva version | SI
| 10381198 | SHBG (GLAE) 200 tests | SI
| 10381461 | Pyrilinks-D 200 tests | SI
| 10381477 | Osteocalcina 200 tests | SI
| 10381450 | Peptido C 200 tests | SI
| 10381455 | Insulin 200 tests | SI
| 10381456 | Insulin 600 tests | SI
| 10381441 | Intact PTH(mono/poly) 200 tests | SI
| 10381442 | Intact PTH(mono/poly) 600 tests | NO
| 10380908 | Ferritin 200 tests | SI
| 10380906 | Ferritin 600 tests | NO
| 10381453 | Micro Albumin 200 tests | SI
| 10487628 | EPO 200 tests | SI
| 10380911 | Folic Acid 200 tests | SI
| 10380912 | Folic Acid 600 tests | NO
| 10380914 | Vitamin B12 200 tests | SI
| 10380913 | Vitamin B12 600 tests | NO
| 10381043 | Proteína C Reactiva 200 tests | SI
| 10381040 | Homocysteine 200 tests | SI
| 10381187 | AFP 200 tests | SI
| 10381184 | AFP 600 tests | NO
| 10380994 | CEA 200 tests | SI
| 10380995 | CEA 600 tests | NO
| 10380986 | PSA 200 tests | SI
| 10380996 | PSA 600 tests | SI
| 10380984 | Free PSA 200 tests | SI
| 10380975 | PAP 200 tests | NO
| 10368607 | ANTI-CCP IgG KIT 200T | SI
| 10380983 | BR-MA (CA 15-3) 200 tests | SI
| 10380972 | OM-MA (CA 125) 200 tests | SI
| 10380988 | GI-MA (CA 19-9) 200 tests | SI
| 10380992 | Beta-2 Microglobulin 200 tests | SI
| 10381539 | Theophylline 200 tests | NO
| 10381550 | Carbamazepine 200 tests | SI
| 10381548 | Phenytoin 200 tests | SI
| 10381544 | Valproic Acid 200 tests | SI
| 10381542 | Phenobarbital 200 tests | SI
| 10381537 | Tobramycin 200 tests | NO
| 10381033 | CK-MB 200 tests | NO
| 10381034 | Troponin I 200 tests | SI
| 10381031 | Myoglobin 200 tests | NO
| 10381036 | NT Pro BNP 200 tests | SI
| 10381041 | Dimero D 200 tests | NO
| 10381452 | IL6 KIT 200 T IMMULITE 2000 | NO
| 10381306 | Hepatitis B surface (HBsAg) 200 tests | SI
| 10381311 | Anti-Hepatitis B core (Anti-HBc) 200 tests | SI
| 10381321 | Anti-HBc IgM 200 tests | SI
| 10381318 | Anti-HBs 200 tests | SI
| 10381323 | Toxoplasma Quantitative IgG 200 tests | SI
| 10381298 | Toxoplasma IgM u-Capture 200 tests | SI
| 10381338 | Rubella Quantitative IgG 200 tests | SI
| 10381305 | Rubella Quantitative IgG 600 tests | NO
| 10381327 | Rubella IgM 200 tests | SI
| 10381333 | Herpes I & II IgG 600 tests | NO
| 10381309 | CMV IgG 200 tests | SI
| 10381320 | CMV IgM 200 tests | SI
| 10488005 | EVB VCA IgM 200 tests | SI
| 10381330 | EVB VCA IgG 200 tests | SI
| 10381307 | EVB EBNA 200 tests | NO
| 10381336 | H. Pylori IgG 200 tests | SI
| 10381335 | H. Pylori IgG 600 tests | NO
| 10380873 | Total IgE 200 tests | SI
| 10380872 | Total IgE 600 tests | SI
| 10380878 | AlaTOP Allergy Screen 200 tests | NO
| 10713448 | TSI kit 200 tests | NO
CONTROLES IMMULITE 2000
| 10282764 | Bio Rad Lipochek Immunoassay Plus (Trilevel 3 x 4 x 5ml) | SI
| 10710159 | Biorad Liquicheck Tumor Marker Control - Trilevel MiniPack (3 x 2ml) | SI
| 10282768 | Biorad Liquicheck Cardiac Markers Plus LT - Nivel 1 (6 x 3 ml) | SI
| 10282769 | Biorad Liquicheck Cardiac Markers Plus LT - Nivel 2 (6 x 3 ml) | SI
| 10282770 | Biorad Liquicheck Cardiac Markers Plus LT - Nivel 3 (6 x 3 ml) | SI
| 10381639 | Thyroglobulin Recovery 2x1ml | SI
| 10385243 | Pyrilinks-D Control Module 2x2ml | SI
| 10385244 | TBG Control Module 2x2ml | SI
| 10385246 | Thyroglobulin Control Module 3x2ml | SI
| 10385316 | C-Reactive Protein Control Module 3x2ml | SI
| 10385317 | PAPP-A Control Module 2x1ml | SI
| 10385337 | LBP Control Module 2x1ml | NO
| 10385349 | Total IgE Control Module 2x2ml | SI
| 10385354 | Intact PTH Control Module 6x2ml | SI
| 10385358 | HCG Control Module 2x2ml | SI
| 10385381 | Autoantibody (TPO - TG) Control Module 2x5ml | SI
| 10385382 | ACTH Control Module 2x2ml | SI
| 10385383 | Control Calcitonin 2x3ml | SI
| 10385384 | EPO Control Module 3x2ml | SI
| 10385385 | Free Beta HCG Control Module 3x1ml | SI
| 10385387 | IGF-1 e IGF-BP3 Control Module 2x2ml | SI
| 10385388 | Albumin Control Module 2x2ml | SI
| 10385389 | Control Interleuquinas Module 2x5ml | NO
| 10385395 | 3rd Generation TSH Control Module (aprox 0,010 mUI/ml) 2x2ml | NO
| 10385397 | Serum Drug Control Module 2x5ml | SI
| 10385587 | ß-2 Microglobulin Control Module 3x1ml | SI
| 10385588 | PAP Control Module 3x2ml | NO
| 10385589 | C-Peptide Control Module 3x2ml | SI
| 10385339 | Control NT PRO BNP | NO
| 10385378 | Control Ferritina | SI
| 10385330 | Control para Dimero D | NO
| 10385331 | Control para Gastrina | SI
| 10385415 | Control Osteocalcina 2x2ml | SI
CONSUMIBLES Y AUXILIARES IMMULITE 2000
| 10385231 | Solución de Lavado 2000 tests | SI
| 10385232 | Sustrato 2000 tests | SI
| 10385206 | Copa de Reacción para Immulite 2K 1000 tests | SI
| 10385207 | Tubos para micromuestras | SI
| 10385209 | Soporte para tubos de micromuestras | SI
| 10387051 | HCG 1x25ml | SI
| 10387054 | Fólico / Vitamina B12 1x25ml | SI
| 10387055 | Homocisteína 1x25ml | NO
| 10387057 | Multidiluyente 1 1x25ml | SI
| 10283031 | MUltidiluyente 2 1x25ml | SI
| 10387058 | MUltidiluyente 2 1x55ml | SI
| 10387060 | PTH 1x25ml | SI
| 10387061 | TSH 1x25ml | SI
| 10387064 | FSH 1x25ml | NO
| 10387067 | OM-MA 1x25ml | NO
| 10387069 | Prolactin 1x25ml | SI
| 10387063 | Estradiol 1x25ml | NO
| 10387073 | Thyroglobulin 1x25ml | SI
| 10387068 | C-Peptide 1x25ml | SI
Menú Alergenos para Immulite 2000
Consumibles
L2KUN6 | 10380875 | 3gAllergy - IgE Especifica x 600
400930-02 | 10291768 | Soporte para Alergenos #001-033 x 33 soportes
| 10282849 | Septum caps x 250
Animales
E1L4 | 10385642 | Caspa y epitelio de gato x 40
E2L4 | 10385648 | Epitelio de perro x 40
E3L4 | 10385650 | caspa de caballo x 40
E4L2 | 10385651 | caspa de vaca x 20
E5L4 | 10385653 | caspa de perro x 40
E6L2 |  | epitelio de conejillo de indias x 20
E7L2 | 10385663 | excremento de paloma x 20
E70L2 | 10385655 | plumas de ganso x 20
E71L2 |  | epitelio de ratón x 20
E72L2 |  | orina de ratón x 20
E73L2 |  | epitelio de rata x 20
E74L2 |  | orina de rata x 20
E75L2 |  | proteínas sericas de rata x 20
E76L2 |  | proteinas sericas de raton x 20
E78L2 |  | plumas de periquito x 20
E80L2 |  | epitelio de cabra x 20
E81L2 |  | epitelio de oveja x 20
E82L4 | 10385667 | epitelio de conejo x 40
E83L2 |  | epitelio de cerdo x 20
E84L2 |  | epitelio de hamster x 20
E85L2 | 10385670 | plumas de pollo x 20
E86L2 | 10385671 | plumas de pato x 20
E87L2 |  | rata x 20
E88L2 |  | ratón x 20
E89L2 |  | plumas de pavo x 20
E91L2 |  | plumas de loro x 20
E201L2 |  | plumas de canario x 20
E209L2 |  | proteinas sericas de pollo x 20
Drogas
C1L2 | 10385620 | penicilina G x 20
C2L2 | 10385624 | penicilina V x 20
C70L2 |  | insulina porcina x 20
C71L2 |  | insulina bovina x 20
C73L2 | 10385627 | insulina humana x 20
C203L2 | 10385621 | ampicilina x 20
C204L2 | 10385622 | amoxicilina x 20
Polvo
H1L4 | 10385940 | polvo doméstico - greer x 40
H2L4 | 10385942 | polvo doméstico - stier x 40
H3L2 |  | polvo doméstico - bencard x 20
H6L2 |  | polvo doméstico - japón x 20
Comida
F1L4 | 10385704 | huevo blanco x 40
F2L4 | 10385784 | leche x 40
F3L4 | 10385815 | bacalao x 40
F4L4 | 10385834 | trigo x 40
F5L2 | 10385841 | centeno x 20
F6L2 | 10385845 | cebada x 20
F7L2 | 10385856 | avena x 20
F8L4 | 10385868 | maíz x 40
F9L2 | 10385877 | arroz x 40
F10L2 | 10385687 | semillas de sesamo x 20
F11L2 |  | trigo negro x 20
F12L2 | 10385690 | arveja x 20
F13L4 | 10385692 | maní x 40
F14L4 | 10385695 | poroto de soja x 40
F15L2 | 10385696 | poroto x 20
F17L4 | 10385699 | avellana x 40
F18L2 |  | nuez de brasil x 20
F20L2 | 10385714 | almendra x 20
F23L2 | 10385739 | cangrejo x 20
F24L4 | 10385745 | camarón x 40
F25L4 | 10385753 | tomate x 40
F26L2 | 10385761 | carne porcina x 20
F27L2 | 10385770 | carne vacuna x 20
F31L4 | 10385797 | zanahoria x 40
F33L4 | 10385803 | naranja x 40
F35L2 | 10385809 | papa x 20
F36L2 | 10385811 | coco x 20
F37L2 | 10385812 | mejillon azul x 20
F40L2 | 10385824 | atun x 20
F41L2 | 10385825 | salmon x 20
F44L2 | 10385827 | frutillas x 20
F45L2 | 10385828 | levadura x 20
F47L2 | 10385829 | ajo x 20
F48L2 | 10385830 | cebolla x 20
F49L4 | 10385832 | manzana x 40
F50L2 |  | caballa x 20
F54L2 |  | batata x 20
F55L2 |  | mijo comun x 20
F56L2 |  | mijo cola de zorro x 20
F61L2 | 10385842 | sardina x 20
F75L4 | 10385847 | yema de huevo x 40
F76L4 | 10385849 | alfa lactalbumina x 40
F77L4 | 10385851 | beta lactalbumina x 40
F78L4 | 10385853 | caseina x 40
F79L4 | 10385855 | gluten x 40
F80L2 | 10385857 | langosta x 20
F81L2 | 10385858 | queso cheddar x 20
F82L2 | 10385859 | queso elaborado con mohos x 20
F83L2 | 10385860 | carne de pollo x 20
F84L2 | 10385861 | kiwi x 20
F85L2 | 10385862 | apio x 20
F86L2 | 10385863 | perejil x 20
F87L2 | 10385864 | melon x 20
F88L2 | 10385865 | cordero x 20
F89L2 | 10385866 | mostaza x 20
F90L2 |  | malta x 20
F91L2 | 10385870 | mango x 20
F92L2 | 10385871 | banana x 20
F93L2 | 10385872 | cacao x 20
F94L2 | 10385873 | pera x 20
F95L2 | 10385874 | durazno x 20
F96L2 | 10385875 | palta x 20
F105L2 | 10385686 | chocolate x 20
F201L2 |  | nuez lisa x 20
F202L2 |  | castaña de caju x 20
F203L2 | 10385707 | pistacho x 20
F204L2 | 10385708 | trucha x 20
F207L2 | 10385711 | almeja x 20
F208L2 | 10385712 | limon x 20
F209L2 | 10385713 | pomelo x 20
F210L2 | 10385715 | anana x 20
F212L2 | 10385716 | seta x 20
F214L2 | 10385718 | espinaca x 20
F215L2 | 10385719 | lechuga x 20
F216L2 | 10385720 | repollo x 20
F217L2 |  | col de bruselas x 20
F218L2 |  | pimenton x 20
F219L2 |  | semilla de hinojo x 20
F221L2 | 10385726 | café x 20
F224L2 |  | semilla de amapola x 20
F225L2 | 10385730 | zapallo x 20
F231L2 |  | leche hervida x 20
F232L2 | 10385732 | ovoalbumina x 20
F233L2 | 10385733 | ovomucoide x 20
F234L2 | 10385735 | vainilla x 20
F235L2 | 10385736 | lenteja x 20
F237L2 | 10385738 | albaricoque x 20
F242L2 | 10385740 | cereza x 20
F244L2 | 10385741 | pepino x 20
F245L2 | 10385742 | huevo x 20
F254L2 |  | platija x 20
F255L2 | 10385748 | ciruela x 20
F256L2 | 10385749 | nuez x 20
F259L2 |  | uva x 20
F260L2 | 10385754 | brocoli x 20
F261L2 | 10385755 | esparrago x 20
F262L2 | 10385756 | berenjena x 20
F263L2 |  | pimienta verde x 20
F269L2 | 10385760 | albahaca x 20
F270L2 | 10385762 | jengibre x 20
F271L2 |  | anis x 20
F272L2 |  | estragon x 20
F273L2 | 10385765 | tomillo x 20
F274L2 |  | mejorana x 20
F277L2 |  | eneldo x 20
F278L2 | 10385768 | hoja de laurel x 20
F279L2 |  | pimienta roja x 20
F280L2 | 10385771 | pimienta negra x 20
F281L2 |  | curry x 20
F283L2 | 10385773 | oregano x 20
F284L2 | 10385774 | carne de pavo x 20
F287L2 |  | poroto rojo arriñonado x 20
F288L2 | 10385776 | arandano x 20
F289L2 |  | datil x 20
F290L2 | 10385778 | ostra x 20
F291L2 | 10385779 | coliflor x 20
F293L2 | 10385780 | papaya x 20
F294L2 |  | pasionaria x 20
F299L2 | 10385782 | castaña x 20
F300L2 |  | poroto pinto x 20
F301L2 |  | caqui x 20
F309L2 | 10385790 | garbanzos x 20
F314L2 |  | caracol x 20
F315L2 | 10385794 | poroto verde x 20
F338L2 | 10385801 | escalope x 20
F401L2 |  | nispero x 20
F402L2 |  | higo x 20
F403L2 | 10385818 | levadura de cervecero x 20
F405L2 |  | menta x 20
F406L2 |  | arugula x 20
| 10386068 | Tilo - Linden x 20
Pastos
G1L2 |  | cesped de primavera x 20
G2L4 | 10385917 | cesped de bermuda x 40
G3L4 | 10385919 | cesped del huerto x 40
G4L2 | 10385920 | cesped de la pradera x 20
G5L4 | 10385922 | cesped perenne de centeno x 40
G6L4 | 10385924 | cesped timothy x 40
G7L2 |  | cesped comun de caña x 20
G8L4 | 10385929 | cesped de junio x 40
G9L2 |  | red top x 20
G10L2 |  | cesped johnson x 20
G11L2 |  | cesped brome x 20
G12L4 |  | cesped de centeno x 40
G13L2 | 10385908 | cesped de terciopelo x 20
G14L2 |  | cesped de avena x 20
G15L2 |  | cesped de trigo x 20
G16L2 |  | cesped cola de zorro de la pradera x 20
G17L2 |  | cesped de bahia x 20
G70L2 |  | cesped de centeno silvestre x 20
G71L2 |  | cesped canario x 20
G202L2 | 10385914 | maiz cultivado x 20
Insectos
I1L4 | 10385950 | veneno de abeja x 40
I2L2 |  | avispon de cara blanca x 20
I3L2 | 10385956 | veneno de avispas amarillas x 20
I4L2 | 10385957 | veneno de abispas de papel x 20
I5L2 |  | avispon amarillo x 20
I6L2 | 10385959 | cucaracha x 20
I8L2 |  | polilla x 20
I70L2 | 10385960 | hormiga x 20
I71L2 | 10385961 | mosquito x 20
I73L2 |  | Chironomus plumosus x 20
I75L2 |  | avispon europeo x 20
I204L2 | 10385952 | tabano x 20
Garrapata
D1L4 | 10385630 | Dermatophagoides pteronyssinus x 40
D2L4 | 10385633 | Dermatophagoides farinae x 20
D3L2 |  | Dermatophagoides microceras x 20
D70L2 | 10385635 | Acarus siro x 20
D71L2 | 10385636 | Lepidoglyphus destructor x 20
D73L2 |  | Glycyphagus domesticus x 20
Mohos
M1L4 | 10386019 | Penicillium notatum x 40
M2L4 | 10386027 | cladosporium herbarum x 40
M3L4 | 10386036 | aspergillus fumigatus x 40
M4L2 |  | mucor racemosus x 20
M5L4 | 10386040 | candida albicans x 40
M6L4 | 10386042 | alternaria tenuis x 40
M7L2 |  | botrytis cinerea x 20
M8L2 |  | helminthosporium halodes x 20
M9L2 |  | fusarium moniliforme x 20
M10L2 |  | stemphylium botryosum x 20
M11L2 | 10386012 | rhizopus nigricans x 20
M12L2 |  | aureobasidium pullulans x 20
M13L2 |  | phoma betae x 20
M14L2 |  | epicoccum purpurascens x 20
M15L2 |  | trichorderma viride x 20
M16L2 | 10386017 | curvularia lunata x 20
M70L2 |  | pityrosporum orbiculare x 20
M202L2 |  | cephalosporium acremonium x 20
M203L2 |  | trichosporom pullulans x 20
M204L2 |  | ulocladium chartarum x 20
M207L2 | 10386023 | aspergillus niger x 20
M208L2 |  | chaetomium globosum x 20
M212L2 | 10386025 | micropolyspora faeni x 20
M300L2 |  | eurotium x 20
M304L2 |  | aspergillus oryzae x 20
M305L2 |  | penicillium brevicompactum x 20
M309L2 | 10386031 | aspergillus terreus x 20
M310L2 | 10386032 | aspergillus nidulans x 20
M311L2 | 10386033 | aspergillus flavus x 20
M312L2 |  | aspergillus clavatus x 20
Ocupacional
K20L2 | 10385980 | lana x 20
K71L2 |  | semilla de ricino x 20
K72L2 |  | isaghula o psyllium x 20
K74L2 |  | seda x 20
K75L2 | 10385990 | isotiocianato TDI x 20
K81L2 |  | ficus x 20
K82L4 | 10385996 | latex x 40
K83L2 |  | semilla de algodón x 20
K84L2 |  | semilla de girasol x 20
K87L2 | 10385999 | alfa amilasa fungal x 20
K76L2 | 10385991 | isotiocianato MDI x 20
K77L2 | 10385992 | isotiocianato HDI x 20
K80L2 | 10385993 | formaldehido x 20
K210L2 |  | anhidrido maleico x 20
K213L2 |  | anhidrido tetracloroftalico x 20
K301L2 | 10385985 | polvo de harina x 20
Otros
O1L2 | 10386049 | algodón x 20
O72L2 | 10386051 | enterotoxina A de S. aureus x 20
O73L2 | 10386052 | enterotoxina B x 20
Parasitos
P1L2 | 10386053 | Ascaris x 20
P4L2 | 10386054 | Anisakis x 20
Arboles
T1L2 | 10386066 | arce x 20
T2L4 |  | aliso x 40
T3L4 | 10386087 | abedul x 40
T4L4 | 10386097 | avellano x 40
T5L2 |  | haya x 20
T6L2 |  | cedro de la montaña x 20
T7L4 | 10386107 | roble x 40
T8L2 | 10386112 | olmo x 20
T9L4 | 10386115 | olivo x 40
T10L2 | 10386056 | nogal x 20
T11L2 | 10386057 | sicomoro x 20
T12L2 | 10386058 | sauce x 20
T14L2 | 10386059 | alamo americano x 20
T15L2 | 10306060 | fresno blanco x 20
T16L2 | 10386061 | pino blanco x 20
T17L4 |  | cedro japones x 40
T18L2 | 10386064 | eucalipto x 20
T19L2 | 10386065 | acacia x 20
T20L2 |  | mesquite, algarroba x 20
T21L2 |  | arbol de corteza de papel x 20
T22L2 |  | nogal liso, nogal morado x 20
T23L4 | 10386079 | cipres italiano x 40
T70L2 |  | morena blanca x 20
T72L2 |  | palmera reina x 20
T73L2 |  | pino autraliano x 20
T77L2 |  | roble mezcla x 20
T80L2 |  | cipres japones x 20
T81L2 |  | aliso japones x 20
T103L2 |  | roble perenne x 20
T209L2 |  | horn beam x 20
T401L2 | 10386088 | arbol de pimienta x 20
T402L2 |  | gomero x 20
T404L2 |  | arbol del paraiso x 20
| 10385631 | Blomia tropicalis x 20
T406L2 |  | palmera x 20
Hierbas
W1L4 | 10386143 | ambrosia comun x 40
W2L2 |  | ambrosia del oeste x 20
W3L2 |  | ambrosia gigante x 20
W4L2 |  | ambrosia falsa x 20
W5L2 |  | artemisa x 20
W6L4 | 10386162 | artemisa vulgaris x 40
W7L2 |  | margarita ojo de buey x 20
W8L2 | 10386166 | diente de leon x 20
W9L4 | 10386168 | platano ingles x 40
W10L2 | 10386131 | quinoa x 20
W11L2 |  | cardo ruso x 20
W12L2 |  | vara de san jose x 20
W13L2 |  | hierba de san guillermo x 20
W14L2 | 10386135 | amaranto rustico x 20
W15L2 |  | scale x 20
W16L2 |  | sauco rustico de pantano x 20
W17L2 |  | arbusto de fuego x 20
W18L2 | 10386139 | acedera x 20
W19L4 | 10386141 | parietaria officinalis x 40
W20L2 |  | ortiga x 20
W21L4 |  | hierba del muro x 40
W75L2 |  | atriplex canescens x 20
W82L2 |  | amaranthus palmeri x 20
W203L2 | 10386144 | pollen de colza x 20
W209L2 | 10386145 | ambrosia mezcla x 20
|  | Paneles | Alergenos incluidos
DP1L4 | 10385640 | Panel Garrapata 1 x 40 | Dermatophagoides pterpnyssinus, farinae, microceras, lepidoglyphus destructor, tyrophagus putrescentiae, glycyphagus domesticus, euroglyphus maynei, blomia tropicales
EP1L4 | 10385677 | Panel Animal 1 x 40 | caspa y epitelio de gato, caspa de caballo, caspa de vaca y caspa de perro
EP2L4 | 10368598 | Panel Animal 2 x 40 | caspa y epitelio de gato, caspa de perro, epitelio de conejilla de indias, rata y ratón
EP70L4 |  | Panel Animal 70 x 40 | epitelio de conejillo de indias, epitelio de conejo, epitelio de hamster, rata y ratón
EP71L4 | 10385681 | Panel Animal 71 x 40 | plumas de ganso, plumas de pollo, plumas de pato, plumas de pavo
EP72L4 |  | Panel Animal 72 x 40 | plumas de periquito, plumas de loro, plumas de canario
FP1L4 | 10385882 | Panel Alimentos 1 x 40 | mani, avellanda, nuez de brasil, almendra, coco
FP2L4 | 10385889 | Panel Alimentos 2 x 40 | bacalao, camaron, mejillon, atun y salmon
FP3L4 | 10385891 | Panel Alimentos 3 x 40 | trigo, avena, maiz, semilla de sesamo, trigo negro
FP5L4 | 10385897 | Panel Alimentos 5 x 40 | huevo blanco, leche, bacalao, trigo, mani, poroto de soja
FP6L4 | 10385899 | Panel Alimentos 6 x 40 | trigo, arroz, semilla de sesamo, trigo negro, poroto soja
FP7L4 |  | Panel Alimentos 7 x 40 | huevo blanco, leche, trigo, arroz, mani, poroto soja
FP8L4 |  | Panel Alimentos 8 x 40 | maiz, semilla de sesamo, trigo negro, cangrejo, camaron
FP13L4 |  | Panel Alimentos 13 x 40 | arveja, poroto, zanahoria, papa
FP15L4 | 10385880 | Panel Alimentos 15 x 40 | naranja, manzana, banana, durazno
FP24L4 |  | Panel Alimentos 24 x 40 | avellana, camaron, kiwi, banana
FP25L4 | 10385884 | Panel Alimentos 25 x 40 | semilla de sesamo, levadura de panadero, ajo, apio
FP26L4 |  | Panel Alimentos 26 x 40 | huevo blanco, leche, mani, mostaza
FP27L4 |  | Panel Alimentos 27 x 40 | bacalao, trigo, poroto soja, avellana
FP28L4 |  | Panel Alimentos 28 x 40 | semilla de sesamo, camaron, carne vacuna, kiwi
FP50L4 | 10385893 | Panel Alimentos 50 x 40 | kiwi, mango, banana, anana
FP51L4 | 10385895 | Panel Alimentos 51 x 40 | tomate, zanahoria, papa, ajo, mostaza
FP73L4 | 10385901 | Panel Alimentos 73 x 40 | carne porcina, carne vacuna, carne de pollo, cordero
GP1L4 | 10385932 | Panel Pastos 1 x 40 | Cesped del huerto, cesped de la pradera, cesped perenne de centeno, cesped timothy, cesped de junio
GP2L4 | 10385934 | Panel Pastos 2 x 40 | Cesped de bermuda, cesped perenne de centeno, cesped timothy, cesped de junio, cesped johnson, cesped de bahia
GP3L4 |  | Panel Pastos 3 x 40 | cesped de primavera, cesped perenne de centeno, cesped timothy, cesped de centeno, cesped de terciopelo
GP4L4 |  | Panel Pastos 4 x 40 | cesped de primavera, cesped perenne de centeno, cesped comun de caña, cesped de centeno, cesped de terciopelo
HP1L4 | 10385948 | Panel Polvo 1 x 40 | dermatophagoides pteronyssinus, dermatophagoides farinae, polvo domestico, cucaracha
IP1L4 | 10385968 | Panel inhalantes 1 x 40 | cesped del huerto, cesped timothy, cedro japones, ambrosia comun, artemisia
IP2L4 |  | Panel inhalantes 2 x 40 | cesped timothy, alternaria tenuis, abedul, artemisia
IP3L4 | 10385970 | Panel inhalantes 3 x 40 | dermatophagoides pteronyssinus, caspa y epitelio de gato, epitelio de perro, aspergillus fumigatus
IP4L4 |  | Panel inhalantes 4 x 40 | dermatophagoides pteronyssinus, caspay epitelio de gato, caspa de caballo, caspa de perro, cucaracha, cladosporium herbarum
IP5L4 |  | Panel inhalantes 5 x 40 | cesped de centeno, cesped de trigo, aspergilus fumigatus, alternaria tenius
IP6L4 | 10385972 | Panel inhalantes 6 x 40 | cesped timothy, cladosporium herbarum, alternaria tenius, abedul, artemisia
IP7L4 | 10385974 | Panel inhalantes 7 x 40 | dermatophagoides pteronyssinus, caspa y epitelio de gato, epitelio de caballo, caspa de perro, epitelio de conejo
IP8L4 | 10385975 | Panel inhalantes 8 x 40 | dermatophagoides pteronyssinus, caspa y epitelio de gato, caspa de perro, cesped timothy, cesped de centeno, cladosporium herbarum, abedul, artemisia
IP9L4 | 10385976 | Panel inhalantes 9 x 40 | caspa y epitelio de gato, caspa de perro, cesped de la pradera, alternaria tenuis, platano ingles
IP10L4 |  | Panel inhalantes 10 x 40 | parietaria judaica, platano ingles, olivo, sicomoro, cipres italiano
KP1L4 |  | Panel Ocupacional 1 x 40 | caspa de caballo, caspa de vaca, plumas de ganso, plumas de pollo
MP1L4 | 10386048 | Panel Mohos 1 x 40 | penicillium notatum, cladosporium herbarum, aspergillus fumigatus, candida albicans, alternaria tenuis
TP1L4 | 10386117 | Panel Arbol 1 x 40 | arce, abedul, roble, olmo, nogal
TP2L4 |  | Panel Arbol 2 x 40 | arce, roble, olmo, alamo americano, nogal
TP3L4 | 10386120 | Panel Arbol 3 x 40 | cedro de la montaña, roble, olmo, alamo americano, algarrobo
TP4L4 | 10386122 | Panel Arbol 4 x 40 | roble, olmo, sicomoro, sauce, alamo americano
TP5L4 |  | Panel Arbol 5 x 40 | aliso, avellana, olmo, sauce, alamo americano
TP6L4 |  | Panel Arbol 6 x 40 | arce, abedul, haya, roble, nogal
TP7L4 | 10386128 | Panel Arbol 7 x 40 | olivo, sauce, pino blanco, eucalipto, acacia, arbol de corteza de papel
TP8L4 | 10368617 | Panel Arbol 8 x 40 | arce, abedul, avellana, roble, sicomoro
TP9L4 |  | Panel Arbol 9 x 40 | aliso, abedul, avellana, roble, sauce
WP1L4 | 10386170 | Panel Hierbas 1 x 40 | ambrosia comun, artemisia vulgaris, platano ingles, quinoa, cardo ruso
WP2L4 | 10386172 | Panel Hierbas 2 x 40 | ambrosia del oeste, artemisia vulgaris, platano ingles, quinoa, scale
WP3L4 |  | Panel Hierbas 3 x 40 | artemisia vulgaris, platano ingles, quinoa, vara de san jose,ortiga
WP5L4 |  | Panel Hierbas 5 x 40 | ambrosia comun, artemisia vulgaris, margarita ojo de buey, diente de leon, vara de san jose
WP6L4 | 10386178 | Panel Hierbas 6 x 40 | platano ingles, quinoa, cardo ruso, acedera
WP7L4 |  | Panel Hierbas 7 x 40 | margarita ojo de buey, diente de leon, platano ingles, quinoa, vara de san jose

## Advia Chemistry
Importación Regular
Familia ADVIA Chemistry - ADVIA 1200 - ADVIA 1800
10309492 | Acido Urico 980 tests 7x140 en ADVIA 1200/ 1800 | SI
10341133 | Acido Urico 4020 tests 6x670 en ADVIA 1200 / 1800 | NO
10377508 | Acido Valproico_2 400 tests 4x100 en ADVIA 1200 / 1800 | NO
10311832 | Albumina 2220 tests 4x555 en ADVIA 1800 | SI
10311832 | Albumina 928 tests 4x232 en ADVIA 1200 | SI
10711840 | Albumina BCP ( Púrpura de Bromocresol ) 3200 tests en ADVIA 1200 /1800 | NO
10318221 | Alpha 1 antitripsina (AAT) 200 tests 2x100 en ADVIA 1200 / 1800 | NO
10309500 | ALT (GPT) 2520 tests 7x360 en ADVIA 1200 / 1800 | SI
10318168 | ALT (GPT) 4020 tests 6x670 en ADVIA 1200 /1800 | NO
10309493 | Amilasa 1225 tests 7x175 en ADVIA 1200 / 1800 | NO
10341130 | Amilasa 2485 tests 7x355 en ADVIA 1200 /1800 | SI
10315178 | Amilasa Pancreática 910 tests 7x130 en ADVIA 1200 / 1800 | NO
10286035 | Amonio 200 tests 2x100 en ADVIA 1200 /1800 | NO
10283014 | Antidepresivos Triciclicos (TCA) 400 tests en ADVIA 1200 /1800 (usa calibradores Toxicologicos Neg , 1,2,3 y 4 ) | NO
10318196 | ApoA1 150 tests 2x 75 en ADVIA 1200 / 1800 | SI
10318207 | Apolipoproteina B (APO B) 170 tests 2x 85 en ADVIA 1200/ 1800 | SI
10318208 | ASO_2 200 tests 2x100 en ADVIA 1200/ 1800 | SI
10309496 | AST (GOT) 2520 tests 7x360 en ADVIA 1200 / 1800 | SI
10341132 | AST (GOT) 4020 tests 6x670 en ADVIA 1200/ 1800 | NO
10335934 | Barbituricos en Suero (sBRB) 400 tests en ADVIA 1200 /1800 | NO
10335939 | Benzodiacepinas en Suero (sBNZ) 400 tests en ADVIA 1200 /1800 | NO
10316610 | Bilirrubina Directa2 520 tests 4x130 en ADVIA 1200/ 1800 | SI
10341115 | Bilirrubina Total 2 1424 tests 4x356 en ADVIA 1200 / 1800 | SI
10341113 | Bili Total_2 2688 tests 4x672 en Advia 1200/1800 | NO
10309053 | C3 200 tests 2x100 en ADVIA 1200 / 1800 | SI
10309054 | C4 200 tests 2x100 en ADVIA 1200 / 1800 | SI
10341124 | Calcio 2 (CPC - Cresolftaleína ) 2700 tests ADVIA 1200 / 1800 | NO
10341116 | Calcio 2 1610 tests 7x230 en ADVIA 1200 | SI
10341116 | Calcio 2 2030 tests 7x290 (2030 tests) en ADVIA 1800 | SI
10309071 | Capacidad total de fijación del hierro (TIBC) 320 tests 4x 80 en ADVIA 1200 | SI
10309071 | Capacidad total de fijación del hierro (TIBC) 400 tests 4x100 en ADVIA 1800 | SI
10376342 | Carbamacepina_2 (CARB) 400 tests 4x100 en ADVIA 1200 / 1800 | NO
10361930 | Cistatina C ( marcador de disfuncion renal ) 400 tests 4x100 Advia 1200 /1800 | NO
10309494 | CK NAC 980 tests 7x140 en ADVIA 1200 / 1800 | SI
10376501 | Colesterol_2 1960 tests 7X280 en ADVIA 1200 | SI
10376501 | Colesterol_2 2450 tests 7x350 en ADVIA / 1800 | SI
10315176 | Colinesterasa 480 tests 4x120 en ADVIA 1800 | SI
10315176 | Colinesterasa 936 tests 4x234 en ADVIA 1200 | SI
10335869 | Creatinina (enzimático_2) 1240 tests ADVIA 1200 / 1800 | SI
10309050 | Creatinina_2 4020 tests 6x670 en ADVIA 1200 / 1800 | SI
10327549 | Digoxina 800 tests 4x200 en ADVIA 1200 / 1800 | NO
10326480 | Dióxido de carbono líquido 1200 tests 4x300 en ADVIA 1200 / 1800 | NO
10379635 | Etanol_2 (ETOH) 600 tests en ADVIA 1200 / 1800 (Usa Calibrador Toxoamonia 10309217) | NO
10311894 | Factor Reumatoideo 200 tests 2x100 en ADVIA 1200 / 1800 | SI
10377505 | Fenobarbital_2 400 tests 4x100 en ADVIA 1200 / 1800 ( Usa Calibrador Durg I 10376770) | NO
10377506 | Fenitoina_2 en ADVIA 400 tests 1200 /1800 ( Usa Calibrador Durg I 10376770) | NO
10309049 | Fosfatasa Alcalina AMP 1540 tests 7x220 en ADVIA 1200 | SI
10309049 | Fosfatasa Alcalina AMP 2450 tests 7x350 (2450 tests) en ADVIA 1800 | SI
10341131 | Fosfatasa Alcalina DEA 1420 tests 4x355 en ADVIA 1800 | SI
10341131 | Fosfatasa Alcalina DEA 1540 tests 7x220 en ADVIA 1200 | SI
10309503 | Fósforo Inorgánico 2506 tests 7x358 en ADVIA 1200 /1800 | SI
10318211 | Fósforo Inorgánico 4020 tests 6x670 en ADVIA 1200 / 1800 | NO
10361941 | Fructosamina 200 tests ADVIA 1800 | NO
10309495 | Gamma-GT 980 tests 7x140 en ADVIA 1200 / 1800 | SI
10335891 | Gluc Hexo 3960 tests ADVIA 1200 / 1800 | NO
10492319 | Glucosa Oxidasa 2200 tests 4x550 en ADVIA 1200 | SI
10314164 | Haptoglobina (HAPT) 540 tests 3x180 en ADVIA 1200 / 1800 | NO
10309507 | HDL-Colesterol Directo 1304 tests 4x326 en ADVIA 1200 /1800 | SI
10379673 | Hemogobin A1c_3 Nueva 160 tests ADVIA 1200 / 1800 con Pretratamiento o Automatizada | Si
10341118 | Hierro 2 1015 tests 7x145 en ADVIA 1200 / 1800 | SI
10309033 | IgA2 375 tests 5x 75 en ADVIA 1200 / 1800 | SI
10309031 | IgG2 360 tests 4x 90 en ADVIA 1200 / 1800 | SI
10309032 | IgM2 360 tests 4x 90 en ADVIA 1200 /1800 | SI
10325776 | Lactato 300 tests 3x100 en ADVIA 1200 / 1800 | SI
10309501 | LDH (L-P) 1918 tests 7x274 en ADVIA 1200 / 1800 | SI
10341128 | LDH (L-P) 4020 tests 6x670 en ADVIA 1200 /1800 | NO
10309502 | LDH (P-L) 1790 tests 5x 358 en ADVIA 1200 / 1800 | SI
10311891 | LDL-Colesterol 752 tests 4x188 en ADVIA 1200 / 1800 | SI
10311896 | Lipasa 640 tests 4x160 en ADVIA 1200 / 1800 | SI
10361942 | Lipoproteína (a) Lp (a) 200 tests 2x100 en ADVIA 1200 / 1800 | NO
10312850 | Litio (LITH) 200 tests 2x100 en ADVIA 1200 / 1800 | SI
10309506 | Magnesio 980 tests 7x140 en ADVIA 1200 /1800 | SI
10318197 | MicroAlbumina (uALB) 420 tests 4x105 en ADVIA 1800 | SI
10318197 | MicroAlbumina (uALB) 520 tests 4x130 en ADVIA 1200 | SI
10361939 | Myoglobina 200 tests 2x 100 en ADVIA 1200 / 1800 | NO
10327381 | Paracetamol (ACET) 600 tests 4x150 en ADVIA 1200 / 1800 | NO
10283129 | PCR_2 1000 tests 4x250 en ADVIA 1200 / 1800 | SI
10318309 | PCRwr Amplio Rango 440 tests 2x220 en ADVIA 1200 / 1800 | SI
10318214 | Prealbumina 200 tests 2x100 en ADVIA 1200 / 1800 | NO
10311878 | Proteínas Totales 2800 tests 4x700 en ADVIA 1200 | SI
10311878 | Proteínas Totales 3400 tests 4x850 en ADVIA 1800 | SI
10335885 | Proteínas Urinarias_2 776 tests 4x194 en ADVIA 1200 / 1800 | SI
10327382 | Salicilato (SAL) 600 tests 4x150 en ADVIA 1200 / 1800 | NO
10377503 | Teofilina_2 400 tests en ADVIA 1200 / 1800 ( Usa Calibrador Durg I 10376770) | NO
10326448 | Tobramicina (TOB) 400 tests 4x100 en ADVIA 1200 /1800 | NO
10318218 | Transferrina 440 tests 4x110 en ADVIA 1200 / 1800 | SI
10335892 | Trigliceridos_2 tests 4x282 en ADVIA 1200 | SI
10335892 | Trigliceridos_2 tests 4x358 en ADVIA 1800 | SI
10309051 | Urea 4020 tests 6x670 en ADVIA 1200 / 1800 | SI
10377509 | Vancomicina_2 (VANC) 200 tests 2x100 en ADVIA 1200 /1800 | NO
Drogas de Abuso Screning en Orina
10378761 | Metabolitos de Cocaína (COCA_2) 760 tests 4x 190 | NO
10378766 | Cannabis_2(THC_2) Marihuana 760 tests 4x 190 | NO
10378735 | Ecstasy (XTC) 760 tests 4x 190 | NO
10378762 | Metadona_2 (MDN_2) 760 tests 4x 190 | NO
10378765 | Propoxifeno_2 (PPX_2) 760 tests 4x 190 | NO
10378764 | Phencyclidine_2 (PCP_2) 760 tests 4x 190 | NO
10378758 | Amphetamines_2 (AMPH_2) 760 tests 4x 190 | NO
10378759 | Barbiturate_2 (BARB_2) 760 tests 4x 190 | NO
10378760 | Benzodiazepine_2 760 tests 4x 190 (BENZ_2) | NO
10378763 | Opiate_2 (OP_2) 760 tests 4x 190 | NO
CALIBRADORES Drogas de Abuso
10445406 | Calibrador DAU Emit Nivel 0 ( Drogas de Abuso en Orina) Todas y Tambien para Extasis como nivel 0 | NO
10445407 | Calibrador DAU Emit Nivel 1 ( Drogas de Abuso en Orina) ( para Amph, Opi) | NO
10445408 | Calibrador DAU Emit Nivel 2 ( Drogas de Abuso en Orina) (Amph, Barb,Benz,THC,Coc,Mdn,Opi,Ppx,PCP) | NO
10445409 | Calibrador DAU Emit Nivel 3 ( Drogas de Abuso en Orina) (Amph, Barb,Benz,THC,Coc,Mdn,Opi,Ppx,PCP) | NO
10445410 | Calibrador DAU Emit Nivel 4 ( Drogas de Abuso en Orina) (Barb,Benz,THC,Coc,Mdn,Opi,Ppx,PCP) | NO
10445411 | Calibrador DAU Emit Nivel 5 ( Drogas de Abuso en Orina) (Amph, Barb,Benz,THC,Coc,Mdn,Opi,Ppx,PCP) | NO
10445480 | Calibrador EXTC , Extasis Emit II Plus Nivel 1 ( EXTC en orina calibrador ) | NO
10445481 | Calibrador EXTC , Extasis Emit II Plus Nivel 2 ( EXTC en orina calibrador ) | NO
10445482 | Calibrador EXTC , Extasis Emit II Plus Nivel 3 ( EXTC en orina calibrador ) | NO
10445483 | Calibrador EXTC , Extasis Emit II Plus Nivel 4 ( EXTC en orina calibrador ) | NO
CALIBRADORES Generales
10311898 | Calibrador Apolipoprotein A1/B 5x0,26gr x1ml | NO
10711842 | Calibrador Albumina BCP ( Púrpura de Bromocresol) | NO
10376487 | Calibrador Cistatina C 2mlx5 niveles | NO
10376770 | Calibradro Drug I ( Fenitoina_2 , Fenobarbital_2; Teofilina_2 ) | NO
10376771 | Calibrador Drug II (Ac. Valp_2 -Vancom_2- Carba_2) | SI
10491408 | Calibrador HbA1c Nueva (para usar con reactivo 10379673 con pretratamiento o automatizada) | SI
10311918 | Calibrador HDL/LDL Chol 3x1ml | SI
10285684 | Calibrador Fructosamina | NO
10311929 | Calibrador Liquido Proteínas Específicas (C3,C4, Prealb,Factor Reunmatoideo, ASO,Alfa 1 Antitrip, Transf ) Nuevo | SI
10285017 | Calibrador Lipoproteina (a) ( Lp(a) | NO
10337470 | Calibrador MicroAlbuminuria 5x2ml | SI
10285682 | Calibrador Myoglobina | NO
10311930 | Calibrador para Proteínas Urinarias 3x5ml | SI
10337471 | Calibrador PCR _2 3x1ml | SI
10313256 | Calibrador PCRwr ( Amplio Rango) | SI
10312279 | Calibrador Química Set Point | SI
10311926 | Calibrador Special Chemistry (Lactato,Lipasa,TIBC, Fosfatasa Acida (Total y no Prostática) | SI
10309217 | Calibrador Toxammonia (Acet,Sal, Ethanol,Ammonia) | NO
10376770 | Calibradro Drug I ( Fenitoina_2 , Teofilina_2 ) | NO
10283015 | Calibrador Toxicologico Serum NEG ( para TCA , serum Barbitiricos y serum Benzodiacepinas ) | NO
10283016 | Calibrador Toxicologico Serum CAL 1 ( para TCA , Barbitiricos y Benzodiacepinas en suero ) | NO
10283017 | Calibrador Toxicologico Serum CAL 2 ( para TCA , serum Barbitiricos y serum Benzodiacepinas ) | NO
10283018 | Calibrador Toxicologico Serum CAL 3 ( para TCA , serum Barbitiricos y serum Benzodiacepinas ) | NO
10283019 | Calibrador Toxicologico Serum CAL 4 ( para TCA , serum Barbitiricos y serum Benzodiacepinas ) | NO
10326423 | Calibrator / Diluent CO2 2x21 ml | NO
CONTROLES ADVIA CHEMISTRY
10338722 | BioRad Lyphochek Assayed Chem 1 | SI
10332619 | BioRad Lyphochek Assayed Chem 2 | SI
10285565 | BioRad Lyphochek Diabetes Nivel 1 6X 0,5 ml | NO
10282722 | BioRad Liquichek Lipids Level 1 | SI
10282723 | BioRad Liquichek Lipids Level 2 | SI
10328005 | BioRad Lyphochek Quant Urine Normal | SI
10330356 | BioRad Lyphochek Quant Urine Anormal | SI
10311921 | Liquichek Immunology Control Level 1 - 6 X 1 ML ( Trae valores para ASO y Cistatina C ) | SI
10311922 | Liquichek Immunology Control Level 1 - 6 X 1 ML (Trae valores para ASO y Cistatina C ) | SI
10311923 | Liquichek Immunology Control Level 1 - 6 X 1 ML( Trae valores para ASO y Cistatina C ) | SI
10470794 | BioRad Liquichek Ethanol/Ammonia 1 | NO
10990133 | BioRad Liquichek Ethanol/Ammonia 2 | NO
10470793 | BioRad Liquichek Ethanol/Ammonia 3 | NO
10282721 | Bio Rad Liquicheck TDM Trilevel | NO
10710177 | Bio Rad Control CSF Nivel 1 ( Liquido cefalorraquideo ) | NO
10710178 | Bio Rad Control CSF Nivel 2 ( Liqudo cefalorraquideo ) | NO
10284683 | Bio Rad Microalbumina Liquicheck Nivel 1 ( Presentación Grande 12X10ml ) | NO
10284684 | Bio Rad Microalbumina Liquicheck Nivel 2 ( Presentación Grande 12X10ml ) | NO
10990280 | Bio Rad Qualitative Urine Toxicology -Control Negativo | NO
10990281 | Bio Rad Qualitative Urine Toxicology - Control Positivo | NO
10311161 | Control para HbA1c con Pretratamiento o Directa | SI
10313975 | Control Orina 1 THC (Canabinoide) | NO
10313976 | Control Orina 2 THC (Canabinoide) | NO
10311161 | Control Normal / Anormal HbA1c | SI
ACCESORIOS LINEA ADVIA
10311853 | Cuvette Wash 1x2000ml | SI
10311854 | Cuvette Conditioner 1x1000ml | SI
10311855 | Incubation Oil 1x2,9litros | SI
10311856 | Rgt Probe Wash 1 5x250ml | SI
10311857 | Rgt Probe Wash 2 5x250 ml | SI
10311858 | Rgt Probe Wash 3 1x500ml | SI
10320912 | Lamp Coolant 1x50ml | SI
10286032 | A1C DENATURANT 920ML | SI
10311849 | ISE Buffer 1x2000ml | SI
10311851 | ISE Detergente 2x100ml | SI
10311903 | ISE Serum Std 2 nivelesx100ml | SI
10311904 | ISE Urine Std 2 nivelesx100ml | SI
10309439 | Electrodos Cl x1 | SI
10309440 | Electrodos K x1 | SI
10309441 | Electrodos Na x1 | SI
10309457 | Contenedor de reactivos de 70 ml (50 x pkg) | SI
10309456 | Contenedor de reactivos de 20 ml (50 x pkg) | SI
10309448 | Lampara Halógena x1 | SI
10309450 | Copa de muestras (200 x pkg) | SI
10327496 | Adaptadores para Sample Cup STT/CTT. Pkg x 50 | SI
ACCESORIOS ADVIA 1200 /1800
10309465 | Electrodo Referencia x1 | SI
10309721 | Cuvette RRV 1x21 ( segmentos de reacción es 1 seg x 21 cubetas ) | SI
10316975 | Adaptadores p/contenedores de 20 a 40 ml (20 x pkg) | SI
10723030 | Adaptadores p/contenedores de 20ml a 70 ml (20 x pkg) | SI
10719152 | Adaptadores p/contenedores de 40ml a 70 ml (20 x pkg) | SI
ACCESORIOS ADVIA 1800
10309465 | Electrodo Referencia x1 | SI
10316975 | Adaptadores p/contenedores de 20 ml (20 x pkg) | SI
10321822 | Cuvette, DTT, 1 X 20 x 1 ( segmentos de dilución es 1 segmento x 20 cubetas) | SI
10321772 | Cuvette, RRV, 1 x 17 x 1 ( segmentos de reacción es 1 segmento de 17 cubetas ) | SI
Note on B146: d02075: Glu,Ure,Crea,AcUco,GOT,GPT,FAL,Ca,Fosf,Mg,Hierro,Bili TyD,GGT,CK,TIBC,Lactico,Litio,LDH,PT,Trig,Alb,Amy,Cloro,Sodio,Pot,
Note on B147: d02075: Glu,Ure,Crea,AcUco,GOT,GPT,FAL,Ca,Fosf,Mg,Hierro,Bili TyD,GGT,CK,TIBC,Lactico,Litio,LDH,PT,Trig,Alb,Amy,Cloro,Sodio,Pot,
Note on B148: Para uso en Hemoglobina Glicosilada de ADVIA y Dimension
Note on B149: d02075: ApoA1,ApoB,ColT,HDLCol,LDLCol,LipoA,PCR,Trig.
Note on B150: d02075: ApoA1,ApoB,ColT,HDLCol,LDLCol,LipoA,PCR,Trig.
Note on B151: d02075: TODO EN ORINA AcUco,Arsenico,Barb,Ca,Cloro,Cobre,CortisolLibr,Crea,Fosf,Gluc,Mg, Merc,Microalb,Plomo,Potas,PT,Urea,Zinc,Catecolaminas libres
Note on B152: d02075: TODO EN ORINA AcUco,Arsenico,Barb,Ca,Cloro,Cobre,CortisolLibr,Crea,Fosf,Gluc,Mg, Merc,Microalb,Plomo,Potas,PT,Urea,Zinc,Catecolaminas libres
Note on B156: d02075: Amonio, etanol
Note on B157: d02075: Amonio, etanol
Note on B158: d02075: Amonio, etanol
Note on B159: d02075: Acetaminof,AcValp,Amikacina,clonazepan,cortisol,ciclosporina,diazepan,digoxina,Carbamac, cloramfenicol,Antidepr triciclicos (TCA),Fenitoína, Fenitoina Libre,fenobarbital,Gentamicina,Imipramina,Lidocaina,Litio,Metrotexato,NAPA, nortriptilina,Primidona,procainamida,propranolol,quinidina,salicilato,t3 libre*,t3 total*,t4 libre*, t4 tota*l,teofilina,tobramicina,TSH*,Vancomicna (*) VERIFICAR VALOR PARA DMINESION

## Dimension
Importación Regular
Dimension RxL Max HM - Dimension Xpand HM - Dimension ExL
10444916 | Acetaminofeno o Paracetamol (ACTM) 80 Tests | NO
10444961 | Ácido Láctico (LA) 120 tests | SI
10464329 | Acido Micofenolico ( MPA) 80 Tests | NO
10444967 | Ácido Úrico (URCA) 480 tests | SI
10444943 | Ácido Valpróico 80 tests | NO
10444975 | Albumina (ALB) 480 tests | SI
10313890 | Alfa1 Glicoprot Acida (AGP) 150 tests ) (Usa Calibrador 10330537) | NO
10475530 | ALTI ( GPT modificada) 240 tests | SI
10444965 | Amilasa (AMY) 240 tests | SI
10711991 | Ammonio (AMM) 120 tests | NO
10317831 | Anti-estreptolisina O ( ASO ) Sentinel 144 tests ( Usa ASO Calibrador 10330145) | NO
10445625 | Apo-Lipoprot A-1 Sentinel 70 tests ( Usa Calibrador 10312783) | NO
10445626 | Apo Lipoprot B ( Apo B) Sentinel 275 tests ( Usa Calibrador 10312783) | NO
10444959 | AST (GOT) 360 tests | SI
10444956 | Bilirrubina Directa (DBI) 320 tests | SI
10444957 | Bilirrubina Total (TBI) 480 tests | SI
10444980 | C3 120 tests | SI
10444981 | C4 120 tests | SI
10444949 | Calcio (CA) 480 tests | SI
10444930 | Carbamazepina 80 tests | NO
10285193 | Ciclosporina Común (CSA) 80 tests | NO
10444934 | Ciclosporina Extendida (CSAE) | NO
10464510 | CK-MB (actividad enzimatica) 120 tests (Linealidad 3-125 U/L ) | SI
10444891 | Colesterol (CHOL) 480 tests | SI
10464511 | Creatina Quinase (CKI) 480 tests | SI
10471520 | Creatinina Enzimatica ( EZCREAT) 336 tests | NO
10872079 | Cratinina ( CRE 2) 480 tests | SI
10444932 | Digitoxina 80 tests | NO
10444941 | Digoxina 80 tests | NO
10444940 | Fenitoina 112 tests | NO
10444933 | Fenobarbital 80 tests | NO
10483771 | Factor Reumatoideo (RF) Sentinel 144 tests ( Usa RF Calibrador 11523D) | NO
10444948 | Fosfatasa Ácida Prostática (ACP) 120 tests | NO
10642445 | Fosfatasa Alcalina ALPI ( ALP modificada) 240 tests | SI
10720277 | Fosforo ( PHOS 2) 480 tests | SI
10444960 | Gamma GT( GGT) 288 tests | SI
10444927 | Gentamicina (GENT) 80 tests | NO
10444971 | Glucosa (GLUC) 1440 tests | SI
10483822 | Hemoglobina Glicosilada Directa ( HA1C Directa) 120 tests - Incluye calibradores | SI
10285412 | Haptoglobina ( HPT) 128 tests ( Usa Calibrador 10330537) | NO
10464332 | HDL Directo Colesterol 240 tests | SI
10444945 | Hierro (IRON) 240 tests | SI
10444982 | IGA 120 tests | SI
10444983 | IGG 120 tests | SI
10444984 | IGM 120 tests | SI
10284483 | LDHI ( Lactato deshidrogenasa ) 480 tests | SI
10444890 | LDL Directo Colesterol 120 tests | SI
10444935 | Lidocaina (LIDO) 80 tests | NO
10460277 | Lipasa (LIPL) 120 tests | SI
10286363 | Lipoproteina (a) ( Lp(a)) 60 tests ( Usa Calibrador 10445637) | NO
10444939 | Litio ( LI ) 80 tests | NO
10444963 | Magnesio (MG 120 tests | SI
10445368 | Metotrexato Syva (MTX) 100 tests- Incluye Calibradores-Requiere Flex vacios | SI
10444977 | Microalbumina (MALB) 80 tests | SI
10444936 | N Acetil Procainamida (NAPA) 80 tests | NO
10444978 | PreAlbumina (PALB ) 120 tests | NO
10444937 | Procainamida (PROC) 80 tests | NO
10444894 | Proteina C Reativa (CRP) 120 tests ( Linealidad 0,2-12 mg/dl) | SI
10444979 | Proteinas Totles (TP) 480 tests | SI
10444964 | Pseudocolinesterasa (PCHE) 120 tests | SI
10444903 | rCRP-Rango extendido o Ultrasensible ( x 120 tests) -Linealidad 0,05-25mg/dl | SI
10444915 | Salicilato (SAL) 120 tests | NO
10464331 | Sirolimus (SIRO) 320 tests | NO
10444908 | T4 120 tests | SI
10700795 | Tacrolimus (TACR) 80 tests
10444942 | Teofilina (THEO) 80 tests | NO
10444944 | TIBC Directo ( Capacidad total de Fijación del Hierro) 240 tests | SI
10444928 | Tobramacina (TOBR) 80 tests | NO
10444906 | Triglicerídeos (TGL) 480 tests | SI
10444985 | Transferrina (TRNF) 120 tests | SI
10444950 | UCFP (Proteinas Líquidos /orina) 80 tests | SI
10444969 | Urea (BUN) 480 tests | SI
10444929 | Vancomicina (VANC) 80 tests | NO
Drogas de Abuso Screning en Orina
10444921 | EXTC Urine Ecstasy Screen Flex 80 tests | NO
10444922 | METH Urine Methadone Screen Flex 80 tests | NO
10444923 | OPI Urine Opiate Screen Flex 80 tests | NO
10444920 | COC Urine Cocaine Metabolite Screen Flex 80 Tests | NO
10444919 | BENZ Urine Benzodiazepine Screen Flex 80 Tests | NO
10444925 | THC Urine Cannabinoids Screen Flex 80 Tests | NO
10444926 | AMPH Urine Amphetamine/Metamphetamine Screen Flex 80 Tests | NO
10444924 | PCP Urine Phencyclidine Screen Flex 80 Tests | NO
10444918 | BARB Urine Barbiturate Screen Flex 80 Tests | NO
10459408 | ETOH/ALC Alcohol Flex 80 Tests | NO
TESTS CON MODULO HM
10444897 | CardioPhase CRP ( CCRP) 120 tests (HM) ( Linealidad 0,1-15 mg/dl ) | NO
10444904 | CKMB-massa 160 tests (HM) ( Linealidad de 0,5-300 ng/ml ) | SI
10444946 | Ferritina FERR 120 tests (HM) | SI
10444888 | FPSA (PSA Libre) 120 tests (HM) | SI
10444910 | FT4 120 tests (HM) | SI
10444913 | HCG Total 120 tests (HM) para uso en RXLMax (Calibrador DC430) | SI
10444914 | LHCG Total (HM) (Unicamente para uso en Xpand Plus-(Calibrador DC 430 ) | NO
10444907 | Miogobina 120 tests (HM) | NO
10444889 | TPSA Total 120 tests (HM) | NO
10444896 | L Troponina I 72 tests (HM) (Unicamente para uso en Xpand Plus-Calibrador RC421C) | NO
10444905 | Troponina I 120 tests (HM) | SI
10444911 | TSH 200 tests ( HM) | SI
10444900 | LPBN N-Terminal Pro-Brain Natriuretic Peptid (uso en Xpand Plus) 72 tests | NO
10444901 | PBNP N-Terminal Pro-Brain Natriuretic Peptid ( uso en RXLMax) 120 tests | NO
TESTS CON MODULO LOCI
10483414 | FT3 LOCI 120 tests | NO
10464523 | FT4L LOCI 120 tests | NO
10464524 | TSH LOCI 200 tests | NO
10464338 | LNTP LOCI ( N terminal Pro Brain Natriuretic Peptide) 72 tests (presentacion Chica) | NO
10464526 | NTP LOCI ( N terminal Pro Brain Natriuretic Peptide) 120 tests ( presentacion grande) | NO
10464525 | TNI LOCI Troponina I ) 144 tests | SI
10463371 | VB12 LOCI ( Vitamina B12) 80 tests | NO
10463370 | FOLA LOCI ( Folato ) 80 tests | NO
ISE
10445058 | IMT- Quik Lyte Na /K/Cl Inderect Sensor ( 1 kit por 4 sensores ) | SI
10453444 | IMT Probe kit | SI
10444878 | Standard A (3x1000ml) Quik Lyte | SI
10444881 | Standard B ( 3x300 ml) Quik Lyte | SI
10445051 | Quik Lyte Flush Solution (3x1,000ml) | SI
10444874 | Sample Diluent Quik Lyte ( 6x500ml) | SI
10444873 | Salt BrIdge Solution ( 3x150 ml) | SI
10445038 | IMT probe cleaner 1x12 bottles | SI
10444872 | Quik Lyte Dilution Check 1 btl 60ml | SI
CALIBRADORES
10444996 | A-LDL Colesterol Calibrador 2x5x1 ml | SI
10445527 | Acido Micofenolico MPA Calibrador | NO
10445026 | rCRP Calibrador ( Rango extendido Linealidad 0,05-25 mg/dl) 2x5x1 ml | SI
10444994 | ACP verifier 2x3x1ml ( para Fosfatasa Acida Prostática ) | NO
10464328 | A-HDL Colesterol Calibrador 2x3x1 ml | SI
10459407 | ETOH /ALC (Alcohol) Calibrador 2x4x1ml | NO
10714028 | ALPI Calibrador | SI
10716280 | CHEM I Calibrador 2X3X2ml (UREA / CA /CREA /CRE2 / EZCREA/ GLUC/ LA /URCA) | SI
10444997 | CHEM II Calibrador 2x3x2ml ( MG /FOSF/ TGL ) | SI
10711580 | CHEM III Calibrador 3X2,5 ml ( para AMM 10711991 ) | NO
10444998 | CHOL Colesterol Calibrador 2x3x1ml | SI
10445001 | Ciclosporina Común (CSA) Calibrador | SI
10445004 | Ciclosporina Extendida (CSAE) Calibrador | NO
10464508 | CKI /MBI Calibrador 2x3x1ml | SI
10445014 | Drug Calibrador 2x5x3 ml ( DGNA /LI /PHNO/ PTN / THEO ) | NO
10445005 | Drug Calibrador II 2x5x5 ml ( ACTM /CARBA/ DGTX / GENT/ LIDO / NAPA / PROC / TOBR /VALP /VANC ) | NO
10444870 | Enzyme Diluent 10x10ml ( Diluyente para enzimas ) | SI
10445016 | Enzyme Verifier 2x3x2ml ( AMY / GGT ) | SI
10476170 | Enzyne II Calibrador ( para ALTI ) | SI
10445019 | HCG y LHCG Calibrador2x5 levels | SI
10445006 | TIBC Directo Calibrador 2x3x1ml | SI
10445010 | IRON ( Hierro ) Calibrador 2x3x1,2ml | SI
10284680 | Enzyme I Calibrador o LDI ( Lactato deshidrogenasa ) Calibrador | SI
10460278 | LIPL Calibrador 2x3x1ml Reemplaza al 10444991 | SI
10445022 | MicroAlbúmina Calibrador 2x5x1 ml | SI
10444987 | PCHE (Colinesterasa ) Calibrador 2x3x2ml | SI
10445025 | Pre ALB Calibrador ( Pre Albumina ) | NO
10444988 | Proteína C Reactiva (CRP) Calibrador 2x5x1 ml ( Linealidad 0,2-12 mg/dl ) | SI
10445028 | Proteinas Espec Calibrador 2x5x1 ml ( para IgG /IgM / IgA / C3 / C4 / Transferrrina ) | SI
10444992 | Salicilato Calibrador 2x3x4 ml (ACA Cal) | NO
10464327 | Sirolimus Calibrador | SI
10445029 | T4 Calibrador 2x5x1 ml | SI
10700796 | Tacrolimus Calibrador 5x2 ml | SI
10445013 | TBI / DBI Calibrador 2x3x1ml | SI
10444993 | TP/ALB Calibrador 2x3x2ml | SI
10444989 | UCFP Calibrator 2x5x4 ml ( Orina y LCR Calibrador ) | SI
10445406 | DAU Emit Calibrador Nivel 0 ( Drogas de Abuso en Orina) Tambien se usa para Extasis como nivel 0 | NO
10445407 | DAU Emit Calibrador Nivel 1 ( Drogas de Abuso en Orina) | NO
10445408 | DAU Emit Calibrador Nivel 2 ( Drogas de Abuso en Orina) | NO
10445409 | DAU Emit Calibrador Nivel 3 ( Drogas de Abuso en Orina) | NO
10445410 | DAU Emit Calibrador Nivel 4 ( Drogas de Abuso en Orina) | NO
10445411 | DAU Emit Calibrador Nivel 5 ( Drogas de Abuso en Orina) | NO
10445480 | EXTC , Extasis Emit II Plus Calibrador Nivel 1 ( EXTC en orina calibrador ) | NO
10445481 | EXTC , Extasis Emit II Plus Calibrador Nivel 2 ( EXTC en orina calibrador ) | NO
10445482 | EXTC , Extasis Emit II Plus Calibrador Nivel 3 ( EXTC en orina calibrador ) | NO
10445483 | EXTC , Extasis Emit II Plus Calibrador Nivel 4 ( EXTC en orina calibrador ) | NO
10330145 | ASO Calibrador 1x 5 ml | NO
10312783 | Apo-A1/Apo-B/HDL Calibrador 3X2 ml | NO
10445637 | Lp(a) Calibrador Set 5x1ml | NO
10330537 | Plasmaproteinas Calibrador 4 x 1 ml ( para Alfa 1 Glicoprot Acida Sentinel) | NO
10285414 | RF Calibrador 5X1ml | NO
CALIBRADORES PARA TESTS DE MODULO HM
10445003 | CCRP Calibrador 2x5 levels (HM) ( Cardiophase CRP ) -Linealidad 0,1-15 mg/dl) | SI
10445017 | FERR Calibrador 2x5 levels (HM) | SI
10445023 | mCKMB calibrador ( de la CKMB Masa ) 2x5 levels (HM) | SI
10445024 | MyoglobinA Calibrador (HM) | NO
10445031 | Total y Free PSA (T/F PSA) calibrador 2x5 levels | NO
10445018 | Thyroid (TSH,fT4 )-Calibrador 2x5 x2 ml (HM) | SI
10444986 | Troponin I calibrador 2x5 levels (HM)para ambas LcTNI y cTNI | SI
10445008 | PBNP y LPBNP Calib 2x5 levels HM ( N-Terminal Pro-Brain Natriuretic Peptid ) | NO
CALIBRADORES PARA TESTS DE MODULO LOCI
10463372 | FOLA/VB12 Calibradr de Folato y Vitamina B12 (2 ml x 5 niv Líquido) | NO
En Proceso | NTP/LNTP Calibrador de N terminal Pro Brain Natriuretic Peptide (1 ml x 5 niv Liquido Frozen) | NO
10484354 | FT4/TSHL/FT3 Calibrador e Free T4 , TSH y Free T3 (2 ml x 5 niv Líquido) | NO
10464336 | TNI Calibrador de Troponina I (2 ml x 5 niv Líquido Frozen) | SI
CONTROLES DIMENSION
10338722 | BioRad Lyphochek Assayed Chem 1 | SI
10332619 | BioRad Lyphochek Assayed Chem 2 | SI
10285565 | BioRad Lyphochek Diabetes Nivel 1 6X 0,5 ml | NO
10282722 | BioRad Liquichek Lipids Level 1 | SI
10282723 | BioRad Liquichek Lipids Level 2 | SI
10328005 | BioRad Lyphochek Quant Urine Normal | SI
10330356 | BioRad Lyphochek Quant Urine Anormal | SI
10460300 | BioRad Lyphochek Immunology Plus | SI
10470794 | BioRad Liquichek Ethanol/Ammonia 1 | NO
10990133 | BioRad Liquichek Ethanol/Ammonia 2 | NO
10470793 | BioRad Liquichek Ethanol/Ammonia 3 | NO
10282721 | Bio Rad Liquicheck TDM Trilevel | NO
10710177 | Bio Rad Control CSF Nivel 1 ( Liquido cefalorraquideo ) | NO
10710178 | Bio Rad Control CSF Nivel 2 ( Liqudo cefalorraquideo ) | NO
10710185 | Bio Rad Control SangreTotal 1 (para Inmunosupresores) | NO
10710186 | Bio Rad Control Sangre Total 2 (para inmunosupresores) | NO
10710187 | Bio Rad Control Sangre Total 3 (para inmunosupresores) | NO
10284683 | Bio Rad Microalbumina Liquicheck Nivel 1 ( Presentación Grande 12X10ml ) | NO
10284684 | Bio Rad Microalbumina Liquicheck Nivel 2 ( Presentación Grande 12X10ml ) | NO
10990280 | Bio Rad Qualitative Urine Toxicology -Control Negativo | NO
10990281 | Bio Rad Qualitative Urine Toxicology - Control Positivo | NO
10471574 | More Control para Tacrolimus /Sirolimus/Ciclosporina | NO
10311161 | Control Normal / Anormal HbA1c | SI
CONSUMIBLES
10445042 | Cuvette Cartridge 12,000/cart | SI
10445039 | Empty FLEX Cart. | SI
10445052 | Chemistry wash 1x1,8L | SI
10445043 | Aliquot wheels 8/bag | SI
10445041 | Sample Cups(1.5ml) 1,000/bag | SI
10445040 | Short Sample Cups 1,000/bag | SI
10481507 | CHK (120 tests/box ( para testeo del fotometro ) | SI
10445050 | Printer Paper XL,RxL,Xpand 4 rolls/97m | SI
10445044 | Vessels - HM 4x250 (Copitas para uso en HM y dilución de las muestras) | SI
10445036 | Probe cleaner 1x500ml ( se usa con el Módulo HM) | SI
10445035 | Sample probe cleaner 1L (se usa con el Módulo HM) | SI
10456305 | Lampara DM source | SI
10455823 | Diafragma Nozzle (3/pack) (Usar en caso de necesidad dado que viene incluido en el cuvette cartridge) | SI
1455793 | Reagent Probe ( 2 agujas x kit) | SI
10457734 | Sample Probe ( 3 agujas x kit) | SI
10445571 | STM ALIQUOT SEGMENT ( para uso con VersaCell ) | SI
10445570 | STM PIPETTE TIPS (para uso con VersaCell) | SI
Note on B196: d02075: Glu,Ure,Crea,AcUco,GOT,GPT,FAL,Ca,Fosf,Mg,Hierro,Bili TyD,GGT,CK,TIBC,Lactico,Litio,LDH,PT,Trig,Alb,Amy,Cloro,Sodio,Pot,
Note on B197: d02075: Glu,Ure,Crea,AcUco,GOT,GPT,FAL,Ca,Fosf,Mg,Hierro,Bili TyD,GGT,CK,TIBC,Lactico,Litio,LDH,PT,Trig,Alb,Amy,Cloro,Sodio,Pot,
Note on B198: Para uso en Hemoglobina Glicosilada de ADVIA y Dimension
Note on B199: d02075: ApoA1,ApoB,ColT,HDLCol,LDLCol,LipoA,PCR,Trig.
Note on B200: d02075: ApoA1,ApoB,ColT,HDLCol,LDLCol,LipoA,PCR,Trig.
Note on B201: d02075: TODO EN ORINA AcUco,Arsenico,Barb,Ca,Cloro,Cobre,CortisolLibr,Crea,Fosf,Gluc,Mg, Merc,Microalb,Plomo,Potas,PT,Urea,Zinc,Catecolaminas libres
Note on B202: d02075: TODO EN ORINA AcUco,Arsenico,Barb,Ca,Cloro,Cobre,CortisolLibr,Crea,Fosf,Gluc,Mg, Merc,Microalb,Plomo,Potas,PT,Urea,Zinc,Catecolaminas libres
Note on B203: d02075: No trae valores para ADVIA si para Dimension AFP,HCG,Perfil de Fertilidad,Perfil Tiroideo,TDM,Ferritina,Litio,IgG,IgM,IgA,IgE,TIBC,
Note on B204: d02075: Amonio, etanol
Note on B205: d02075: Amonio, etanol
Note on B206: d02075: Amonio, etanol
Note on B207: d02075: Acetaminof,AcValp,Amikacina,clonazepan,cortisol,ciclosporina,diazepan,digoxina,Carbamac, cloramfenicol,Antidepr triciclicos (TCA),Fenitoína, Fenitoina Libre,fenobarbital,Gentamicina,Imipramina,Lidocaina,Litio,Metrotexato,NAPA, nortriptilina,Primidona,procainamida,propranolol,quinidina,salicilato,t3 libre*,t3 total*,t4 libre*, t4 tota*l,teofilina,tobramicina,TSH*,Vancomicna (*) VERIFICAR VALOR PARA DMINESION

## Hemostasia
Importación Regular
Hemostasia
KPTT
B4218-1 | 10445709 | Actin 10x2ml | SI
B4218-2 | 10445711 | Actin 10x10ml | NO
B4219-1 | 10445713 | Actin Fsl 10x2ml | SI
B4219-2 | 10445714 | Actin Fsl 10x10ml | NO
OQGS295 | 10446066 | Pathromtin Sl 10x 5ml | SI
OQGS355 | 10446067 | Pathromtin Sl Kit 20x5ml | NO
PT
OUHP295 | 10446442 | Thromborel-S 10xf.4ml | SI
OUHP495 | 10446445 | Thromborel-S 10xf.10ml | SI
B4212-50 | 10445706 | Innovin 10xf.10ml | SI
B4212-100 | 10445704 | Innovin 12xf.20ml | NO
Fibrinogeno
OWZG19 | 10446689 | Multifibren U 10xf.2ml | SI
OWZG23 | 10446691 | Multifibren U 10xf.5ml | SI
B4233-25 | 10445720 | Thrombin Reagent 10xf.1ml | SI
B4233-27 | 10445721 | Thrombin Reagent 10xf.5ml | NO
Tiempo De Trombina
OWHM135 | 10446598 | Test-Thrombin-Reagent Kit 10x5 | SI
OUOV21 | 10446463 | Batroxobin Reagent 2x5ml | NO
OWNA115 | 10446636 | Bc-Thrombin-Reagent kit 10x5ml | NO
Factores
OSGR135 | 10446330 | Factor II 3xf.1ml | SI
OTXX175 | 10446414 | Factor IX 8xf.1ml | SI
OTXV135 | 10446407 | Factor VII 3xf.1ml | SI
ORSM195 | 10446269 | Factor V 8 X F.1ml | SI
OSDG135 | 10446318 | Factor XII.3xf.1ml | SI
OSDF135 | 10446316 | Factor XI.3xf.1ml | SI
OTXY135 | 10446415 | Factor X 3xf.1ml | SI
OTXW175 | 10446411 | Factor VIII.8xf.1ml | SI
OWSU115 | 10446652 | Berichrom Factor XIII Testkit 3x5 | SI
B4238-40 | 10445729 | Factor VIII cromogenico | NO
Antithrombina III
OWWR175 | 10446673 | Berichrom Attrombina III 6x5ml | SI
Protein C / S - Determination -
OUVV155 | 10446499 | Berichrom-Protein C kit 3x10ml | SI
OQLS135 | 10446101 | Proc Global Kit 4x2ml | SI
OQYG115 | 10446185 | Protein C-Reagent kit 1x3ml | NO
OPAP035 | 10445968 | Protein S Ac kit 6x1ml | SI
OPBC032 | 10445977 | ProC® Ac R (5 x 2 ml) Resistencia a la Proteína C | SI
OPGL035 | 10446029 | INNOVANCE Prot S libre | SI
Von Willebrand
OPHL03 | 10487040 | Innovance von- Willebrand F Ac (Método automatizado) | NO
OUBD375 | 10446425 | Bc-Von-Willebrand-Reag.5xf.4ml (Método Automatizadol) Ristocetina | SI
OPAB032 | 10445967 | Von Willebrand Antigen (Vwf Ag) 4x2ml | SI
OUBD235 | 10446423 | Von-Willebrand-Reag. 5xf.2ml (Método Manual) Ristocetina | SI
Pruebas Especiales
OWOA155 | 10446642 | Berichrom Pai Kit 1x2ml | NO
OUBU155 | 10446427 | Berichrom-A2-Antiplasmin Kit 3x5ml | NO
OWLD115 | 10446620 | Berichrom-Heparin kit 3x10ml | SI
OVIA15 | 10446446 | Berichrom C1 Inhibidor | SI
OUCA175 | 10446431 | Berichrom-Plasminogen Kit 3x5ml | NO
OPGA03 | 10446023 | Innovance ETP (Potencial de Trombina Endógena) solo para BCS y BCSXP | NO
Dímero D
OPBP035 | 10445979 | INNOVANCE D-Dimer 150 tests/kit | SI
B4233-60 | 10445722 | Dimertest Latex 60 tests/kit | NO
Lupus
OQGP172 | 10446063 | LA 1 10 X 2 Ml | SI
OQGR132 | 10446064 | LA 2 10 X 1 Ml | SI
CONSUMIBLES Y REACTIVOS AUXILIARES
OQAA335 | 10446032 | Reactivo Imidazol 6x15ml | NO
OQAB455 | 10446033 | Kaolin-Suspension 50ml | SI
ORHO375 | 10446232 | Cloruro Calcio 10x 15ml (Rvo Adicional APTT) | SI
OVKG032 | 10446530 | Cubetas Con Mezcladores BFT II 5x100 | SI
OVKH15 | 10446532 | Papel Para BFT II ( por 10 rollos ) | SI
| 10446708 | Papel térmico p CA-500 x 10 Un | SI
964-0631-3 | 10445689 | CA Clean I p CA 560 / CA 660 / CA 1500 / CA CS2000i (1 x 50ml) | SI
974-0581-0 | 10445690 | CA Clean II p CA 660/ CA1500/ CS2000i (1 x 5 Lts) | SI
974-0771-3 | 10445691 | Sample Plates SAP 400A p CA 1500 (2500 pocillos) | SI
GCL-100A | 10445861 | CA CAL S (p calibracion del sistema òptico del CA 560 y CA660) | SI
904-07219 | 10445687 | Tubos de Reaccion CA 560 / CA 660 /CA 1500 (3 X 1000) | SI
064-1481-0 | 10488059 | Tubos de reacción p CS2000i (x 3000) | SI
OVFR032 | 10446513 | Cubetas 6x10 rotores (32 pocillos)/ BCT-BFA | SI
OWZC395 | 10446684 | Solución De Lavado p BCT / BCS / BCSXP (10x15ml) | SI
OVIS312 | 10446526 | Copas 4ml p BCT / BCS / BCSXP (x 500) | NO
424-1160-8 | 10449436 | Copas cónica 4 ml p CA560 / CA660 /CA1500/CS2000i (x100) | NO
OVIP152 | 10446524 | Rotor para BCS / BCS XP 126unid | SI
OQSC115 | 10446125 | Bc Validation kit p BCS / BCSXP | NO
181604 | 10482450 | Terralin Protect - Large Bottle p BCS / BCSXP / CA560 / CA 660 | NO
OVKE492 | 10446529 | Bc-Bottle Kit 45x15ml (frascos para reactivo) p BCSXP | NO
B4265-37 | 10873440 | Ca System Buffer 8x250ml (Buffer veronal) NUEVO | SI
OQUB19 | 10446135 | Cleaner SCS p BCS/BCSXP
541-1352-1 | 10470110 | Push vial 10 (GW5 frascos para reactivos) CA560/CA660/CA1500/CS2000i | NO
D01317714 | 10455382 | SLD Vial Assy 10/Pack (frasco para buffer) CA560/CA660 | NO
AX008688 | 10709524 | SLD Mini Copas p CS2000i (500 uni) | NO
064-1041-9 | 10481761 | Cubetas para Buffer p CS2000i (144 uni) | NO
CALIBRADORES
ORKL 17 | 10446238 | Standard-Human-Plasma 10xf.1ml | SI
OQVK115 | 10446148 | Fibrinogen Calibrator Kit | SI
OPAT035 | 10445969 | Pt - Multi Calibrator niveles 1 - 6 | NO
OPCC035 | 10445989 | Berichrom Heparin UF Calibrator | NO
OPCH035 | 10445987 | Berichrom Heparin LMW Calibrator | NO
OPGE03 | 10446024 | Innovance ETP standard | NO
CONTROLES HEMOSTASIA
OUPZ175 | 10446471 | Plasma Control P 10x F.1ml | SI
ORKE415 | 10446234 | Plasma Control N (H) 10xf.1ml | SI
OPDY 09 | 10446005 | INNOVANCE D-Dimer Controls | SI
OPBY035 | 10445985 | Berichrom Heparin UF Control 1 | NO
OPBZ035 | 10445986 | Berichrom Heparin UF Control 2 | NO
OPCD035 | 10445990 | Berichrom Heparin LMW Control 1 | NO
OPCB035 | 10445988 | Berichrom Heparin LMW Control 2 | NO
OQKE175 | 10446096 | Proc-Plasma Control 6xf.1ml | NO
291070 | 10445601 | CI-TROL 1 (E) 10XF.1ML | NO
291071 | 10445602 | CI-TROL 2 (E) 10XF.1ML | NO
291072 | 10445603 | CI-TROL 3 (E) 10XF.1ML | NO
PFA100-PFA200
B4170 20 | 10445696 | Colageno/Epinefrina cartuchos X 20 determ | NO
B4170-21 | 10445698 | Colageno/ADP cartuchos x 20 determ | NO
B4170-22A | 10445700 | Innovance P2Y cartuchos x 20 determ | NO
B4170-50 | 10445701 | Trigger Solution/solución activadora 3 x 11 ml | NO
CONSUMIBLES/ACCESORIOS
B4170-70 | 10459383 | PFA-100® Cassettes 1 x 5 | NO
B4170-71 | 10459384 | PFA-100® Printer Paper 10 rolls | NO
OPJG03 | 10481748 | PFA200 Printer Paper NUEVO | NO
B4170-72 | 10459385 | PFA-100® Printer Ribbon 1 pieza | NO
B4170-73 | 10459386 | PFA-100® Cleaning Pads 1 x 35 | NO
B4170-74 | 10445702 | PFA-100® Priming Cartridges 10 cartridges | NO
B4170-75 | 10445703 | PFA-100® Vacuum Test Cups 35 cups | NO
B4170-77 | 10459387 | PFA-100® O-Ring Service Tool 1 unit | NO
B4170-78 | 10459388 | PFA-100® Spare O-Rings 1 x 2 piezas | NO
Note on C63: z002s69w: Solo corre en BCS y BCSXP
Note on C107: z002s69w: liofilizado

## Hematologia
Importación Regular
Hemostasia
ADVIA 120 / 2120
8062T01362052 | 10312269 | ADVIA CBC TIMEPAC | SI
8062T01362152 | 10312270 | ADVIA DIFF TIMEPAC FORM | SI
8062T01362251 | 10330709 | ADVIA AUTO RETIC 1X820 | SI
8062T01362301 | 10312272 | ADVIA SHEATH/RINSE 1X20 | SI
| 10285021 | ADVIA EZ Wash (2 x 1620 ml) NUEVO | SI
8062T01363354 | 10312275 | ADVIA PEROX SHEATH | SI
8062T03368552 | 10312285 | ADVIA Set Point Calibrador (2 x 6,1 ml) | NO
ADVIA 2120
T013626-52 | 10341169 | ADVIA 2120 CBC TIMEPAC Hb CN Free | SI
CONTROLES ADVIA 120 / 2120
8062T03441654 | 10316217 | ADVIA 120 Control Normal (3 en1)Test Point Normal 4x4ml | NO
T033686F54 | 10330063 | ADVIA 120 Control (3 en1) Test Point Abnormal 1(bajo) 4x4ml | NO
T033688F54 | 10318905 | ADVIA 120 Control (3 en1) Test Point Abnormal 2(alto) 4x4ml | NO
ADVIA 360
| 11170848 | ADVIA 360 3P Lyse - Hematology - RGT - 900 Tests (1L) | NO
| 11170849 | ADVIA 360 Cleaner Hematology Cons 1L | NO
CONTROLES ADVIA 360
| 11170853 | ADVIA 360 Hematology Control 3-Level 6x3mL | NO
ADVIA 560
| 11170846 | ADVIA 560 Lyse 5P Hematology RGT 900 Tests (5L) | NO
| 11170847 | ADVIA 560 Lyse-2 Hematology Cons 1L | NO
| 11170851 | ADVIA 560 hard clean Hematology Cons 100mL | NO
CONTROLES ADVIA 560
| 11170854 | ADVIA 560 Hematology Control 3-Level 6x3mL | NO
Advia 360 / 560 - Insumos compartidos
| 11170845 | ADVIA 360/560 Diluent Hematoloy Cons 20L | NO
| 11170852 | ADVIA 360/560 Calibrator Hematology 2x3mL | NO

## Plasma Proteinas
Importación Regular
PLASMA PROTEINAS
BN ProSpec
OPBU05 | 10445984 | N Látex ASL 4x3,5ml | SI
OQNM13 | 10446116 | N Látex Cistatina C | NO
OPCE05 | 10445992 | N Látex RF 4x4ml | SI
OQMP11 | 10446108 | N látex SAA | NO
OUIF09 | 10446452 | N-Anti - Prealbúmina 2ml | NO
OSAZ09 | 10446311 | N-Anti a1 - Antitripsina 2ml | NO
OSAZ15 | 10446312 | N-Anti a1 - Antitripsina 5ml | NO
OWLA11 | 10446619 | N-Anti a1 - Microglobulina | NO
OSAW09 | 10446306 | N-Anti a1-Glicoproteína ácida 2ml | NO
OSAW15 | 10446307 | N-Anti a1-Glicoproteína ácida 5ml | NO
OSAM09 | 10446284 | N-Anti a2-Macroglobulina 2ml | NO
OSAM15 | 10446285 | N-Anti a2-Macroglobulina 5ml | NO
OSAL11 | 10446282 | N-Anti-Albúmina 2ml | SI
OSAL15 | 10446283 | N-Anti-Albúmina 5ml | NO
OSAY09 | 10446310 | N-Anti-Antitrombina lll 2ml | NO
OQBA09 | 10446037 | N-Anti-Apo All | NO
OQDL09 | 10446045 | N-Anti-Apo E | NO
OUED09 | 10446437 | N-Anti-Apolipoproteína Al 2ml | NO
OUED15 | 10446438 | N-Anti-Apolipoproteína Al 5ml | NO
OSAN15 | 10446287 | N-Anti-Apolipoproteína B 5ml | NO
OSAN09 | 10446286 | N-Anti-Apolipoproteína B 2ml | NO
OQEY09 | 10446049 | N-Anti-C1 Inhibidor 2ml | NO
OSAP09 | 10446290 | N-Anti-C3c 2ml | SI
OSAP15 | 10446291 | N-Anti-C3c 5ml | NO
OSAO09 | 10446288 | N-Anti-C4 2ml | SI
OSAO15 | 10446289 | N-Anti-C4 5ml | NO
OUIE09 | 10446451 | N-Anti-Ceruloplasmina 2ml | NO
OSCA09 | 10446313 | N-Anti-Fibrinógeno 2ml | NO
OUND09 | 10446462 | N-Anti-Fibronectina 2ml | NO
OSAV09 | 10446304 | N-Anti-Haptoglobina 2ml | NO
OSAV15 | 10446305 | N-Anti-Haptoglobina 5ml | NO
OUVN09 | 10446493 | N-Anti-Hemopexina 2ml | NO
OWHG09 | 10446594 | N-Anti-Ig/Cadena-L-Tipo Kappa 2ml | SI
OWHH09 | 10446596 | N-Anti-Ig/Cadena-L-Tipo Lambda 2ml | SI
OSAR09 | 10446292 | N-Anti-IgA 2ml | SI
OSAR15 | 10446294 | N-Anti-IgA 5ml | NO
OSAS09 | 10446296 | N-Anti-IgG 2ml | SI
OSAS15 | 10446298 | N-Anti-IgG 5ml | NO
OQXI09 | 10446168 | N-Anti-IgG1 1,5ml | NO
OQXK09 | 10446169 | N-Anti-IgG2 1,5ml | NO
OPAV032 | 10445972 | N-Latex IgG3 NUEVO | NO
OPAU032 | 10445971 | N-Latex IgG4 NUEVO | NO
OSAT09 | 10446300 | N-Anti-IgM 2ml | SI
OSAT15 | 10446302 | N-Anti-IgM 5ml | NO
OSCB09 | 10446314 | N-Anti-Plasminógeno 2ml | NO
OUVO09 | 10446494 | N-Anti-RBP 2ml | NO
OSAX09 | 10446308 | N-Anti-Transferrina 2ml | SI
OSAX15 | 10446309 | N-Anti-Transferrina 5ml | NO
OQIY13 | 10446090 | N-Látex Cardio Phase hs CRP 3x2ml | SI
POCS03 | 10445996 | N-Látex CDT | NO
OQTH11 | 10446130 | N-Latex Ferritina | SI
OPAX03 | 10445973 | N-Látex Homocisteina (HCY) | NO
OQTG15 | 10446129 | N-Latex IgE mono 3x3ml | NO
OQTC11 | 10446127 | N-Latex sTfR 3x2ml | NO
OWTI11 | 10446658 | N-Reactivo ADNasa B | NO
OQWU15 | 10446161 | N-Reactivo Latex Beta 2 Microglobulina | SI
NCMH11 | 10445951 | N-Reactivo Látex CIC | NO
OQAI11 | 10446036 | N-Reactivo Látex IgA | SI
OQAC11 | 10446035 | N-Reactivo Látex IgM | SI
OQHL11 | 10446069 | N-Reactivo Látex Lp(a) 3x2ml | SI
OWIA11 | 10446607 | N-Reactivo Látex Mioglobina | NO
OPJA03 | 10482437 | N LATEX FLC KAPPA (3 x 37 test) | NO
OPJB03 | 10482438 | N LATEX FLC LAMBDA (3 x 37 test) | NO
CALIBRADORES
OQCV05 | 10446038 | N-Standard Lp(a) 3x0,5ml | SI
OQIM13 | 10446073 | N-Standard de proteínas SL 3x1ml | SI
OQKZ13 | 10446099 | N-Standard de Reumatología SL 3x1ml | SI
OQLV07 | 10446102 | N-Standard de proteínas urinarias UY 3x0,5ml | SI
OUID13 | 10446449 | N-Standard proteínas PY 3x1ml | SI
OUPG07 | 10446465 | N/T Standard Apolipoproteínas 3x0,5ml | NO
OPJD03 | 10482440 | N FLC STANDARD SL (3x1ml) | NO
CONSUMIBLES
OQTD11 | 10446128 | R. Suplementario L 6x0,5ml | SI
OUMS65* | 10446455 | N- Buffer de Reacción para BNII y BNPS 5 lts | SI
OUMT65 | 10446457 | N-Diluyente para BNII y BNPS 5 lts | SI
OUMU15 | 10446458 | N-Reactivo suplementario/precipitación 5ml | SI
OQUB19 | 10446135 | Cleaner SCS 6x5ml para BNII y BNProspec | SI
OVIC11 | 10446521 | Cubetas para dilución BNII y BN ProsPec 1100x6 | SI
OVLC21 | 10446534 | Tapa de evaporación de BNProspec-Control 1x50 | SI
OVLC31 | 10446535 | Tapa de evaporación de BNProspec-Reactivos 1x100 | SI
OVLK31 | 10446539 | Cubetas lectura BNProSpec 300x9 | SI
OPJC03 | 10482439 | N FLC SUPPLEMENTARY REAGENT 3x2,5ml | NO
CONTROLES
OQIN13 | 10446076 | N/T Control proteínasSL ( Sericas Liquído) , Nivel bajo 3x1,0ml | SI
OQIO13 | 10446082 | N/T Control proteínas SL,Nivel Medio 3x1ml | SI
OQIP13 | 10446086 | N/T Control proteínas SL, Nivel Alto 3x1ml | SI
OQLW13 | 10446104 | N/T Control de proteínas LC 3x1ml | NO
OQCW05 | 10446040 | N Suero Control Lp SY 3x0,5ml | NO
OQDB13 | 10446041 | N Suero Control Reuma SL/1 3x1ml | SI
OQDC13 | 10446043 | N/T-Suero control Reuma SL/2 3x1ml | SI
OUPH07 | 10446468 | N-Suero Control de Apolipoproteínas 3x0,5ml | NO
OWSY13 | 10446655 | N/T Suero Control Proteínas PY 3x1ml | SI
OPJE03 | 10482441 | N FLC CONTROL SL1 (3x1ml ) NUEVO | NO
OPJF03 | 10482442 | N FLC CONTROL SL2 (3x1ml) NUEVO | NO

## Enf. Infecciosas - Serodia
Importación Regular
SERODIA
80633149 | 10338700 | Serodia Chagas 5 x 20 | SI
80630658 | 10333430 | Serodia HIV 1+2 5 x 20 | SI
80631426 | 10331173 | Serodia HIV 5 x 20 | SI
80634595 | 10339422 | Serodia HTLV 5 x 20 | SI
80636384 | 10311536 | Serodia TP - PA 5 x 20 | SI
80638722 | 10339482 | Fastec 401 Tips / Bulk | SI
80638739A | 10340215 | Fastec 401 Tips C/ Rack x 10 | NO
80631433 | 10341268 | Micro Plate U x 5 Unidades | SI
ENFERMEDADES INFECCIOSAS (IDD)
OWBR15 | 10446582 | Enzygnost anti Rubeola G 2x48 | NO
OWBO15 | 10446585 | Enzygnost anti Rubeola M 2x48 | NO
OWLN15 | 10446622 | Enzygnost anti Sarampión IgG 2x48 | NO
OWLI 15 | 10446621 | Enzygnost anti Sarampión IgM 2x48 | NO
OUVP17 | 10446495 | Reactivos complementarios para Enzygnost/TMB | NO

## Molecular
Importación Regular
BU Molecular
GENOTIPIFICACIÓN VIRUS HEPATITIS C (LIPA)
10325050 | VERSANT HCV AMPLIFICATION 2.0 | NO
10325051 | VERSANT HCV CONTROL 2.0 KIT | NO
10325052 | VERSANT HCV GENOTYPE 2.0 | NO

## Orinas
Importación Regular
UROANÁLISIS / Clinitek ATLAS / Clinitek NOVUS
| 10313959 | Multistix 10-SG (10 parámetros) x 100 tiras | SI
| 10316287 | Chek-Stix Combo: 25 positivas/25 negativas (Orina Control) | SI
80632083 | 10318844 | Clinitek Microglobulina/Creatinina x 25 tiras Para ser leido en Clinitek Status + | SI
80635017 | 10311120 | Clinitek Atlas Reagent Pack (1r X 490 Tests) | SI
80635018A | 10311123 | Clinitek Atlas Calibration Kit (Cal 1,2,3 y 4) 4 x 230ml | SI
80635007A | 10311113 | Clinitek Atlas Rinse Additive (4 x 26 ml) | SI
| 10634643 | CLINITEK NOVUS 10 CASSETTE x 450 Tests | Si
| 10634644 | CLINITEK NOVUS PRO 12 CASSETTE 450 (con relación de creatinina y proteinas) | No
| 10697753 | CLINITEK NOVUS CALIBRATOR | Si
| 10697754 | CLINITEK NOVUS RINSE ADDITIVE | si
CONTROLES Clinitek ATLAS y Clinitek NOVUS
80635019 | 10311124 | Clinitek Atlas Control Positivo X 25 tiras | SI
80635037 | 10311135 | Clinitek Atlas Control Negativo X 25 tiras | SI
SISTEMA DCA 2000 / DCA VANTAGE
80635035 | 10311134 | DCA 2000 HbA1c x 10 det. ind (6 minutos). | SI
80636011 | 10311480 | DCA 2000 Microalbúmina/Creatinina x 10 det. ind.(7’) | SI
| 10314709 | Paper Printer Rolls 1 box (5 rollos) p/ DCA Vantage ( nuevo) | NO
DCA 2000/ DCA VANTAGE
80635068 | 10311161 | Control Normal / Anormal HbA1c | SI
80636012 | 10325406 | Control Microalbúmina / Creatinina | SI

## Cuidados Criticos
Importación Regular
CUIDADOS CRITICOS
Rapidlab 248:pH/pO2/pCO2 Barómetro incorporado
8062473496 | 10310996 | Buffer Pack 6.838 buffer: 4 botellas x 90 ml + 7.382 buffer: 4 botellas x 370 ml | SI
8062473497 | 10310997 | Wash & CD PackDesprot.& conditioning pack x 4 | SI
8062673358 | 10311662 | Sample & Reagent Pump Tubing Kit (1Unidad) | SI
8062673251 | 10311627 | Probe & Tubing Kit - Aguja y Tubuladura (1Unidad) | SI
8062673253 | 10311628 | Probe & Housing Kit - Módulo Toma de Muestra (1Unidad) | SI
8062673250 | 10317482 | Bottle Tubing Kit (1Unidad) | SI
8062477434 | 10311037 | Tubo de Gas 1 de Calibración 1 Tubo (para todos) (5%CO2-12%O2, Balance N2) | SI
8062477438 | 10311038 | Tubo de Gas 2 Slope 1Tubo (248 y Serie 800) (10%CO2- 0%O2, Blance N2) | SI
8062127191 | 10310422 | Regulador doble para ambos tubos 1 unidad | SI
Rapidlab 348:pH/pO2/pCO2/Na/K/Ca/Cl/Hcto Barómetro incorporado
8062104227 | 10309757 | Buffer Pack 4 botellas | SI
8062104226 | 10309756 | Wash & CD Pack 4 botellas | SI
8062105670 | 10309776 | Hcto. Slope 10 ampollas x 2 ml. | SI
8062106042 | 10309783 | Hct Electrode (1 unidad) | SI
8062105673 | 10309779 | Reagent & Sample Pump Tubing Kit (1 Unidad) | SI
8062105672 | 10309778 | Bottle Tubing Kit (1 Unidad) | SI
8062105070 | 10309768 | Gas Cartridge Pack (Calibration & Slope) x 2 Tubos | SI
8062107275 | 10309797 | Probe & Tubing kit Aguja y Tubuladura (1 Unidad) | SI
8062673253 | 10311628 | Probe & Housing kit - Módulo Toma de Muestra (1 Unidad) | SI
8062673373 | 10324749 | Probe Protectors - Arandela Protectora toma de Muestra (Pack x 10) | SI
8062570097 | 10311387 | High G/L Solution 10 ampollas de 2.5 ml. | SI
| 10329817 | QUICK ADAPTER 800S 100PK | NO
Rapidlab SERIE 1200: ph/ pO2/ pCo2/ Na/ K/ Ca/ Cl/ Gluc/ Lac/ Co-ox/ pAt
806203909458 | 10319811 | Reagent Cartridge (1 unidad) | SI
806203913056 | 10319817 | Wash Cartridge ( 2 unidades) | SI
806206324604 | 10324165 | Co-ox sampler chamber RL 1200 | SI
806206641057 | 10324896 | Waste Bottle (4 unidad) RL1200 | NO
8062101404 | 10315772 | Printer paper RL 1200 | SI
806200792177 | 10314089 | Printer paper spindle RL1200 | SI
806209340910 | 10329857 | Printer paper cover Assy RL1200 | SI
806206646415 | 10324912 | Reagent Manifold 1260 | NO
806206643971 | 10324903 | Reagent Manifold 1265 | NO
806206645048 | 10324907 | Tubing kit RL1200 | SI
806203705089 | 10311373 | Co-ox pump tube kit RL1200 | SI
806206451843 | 10324388 | Reference electrode refill cassette | SI
806205065575 | 10321856 | pO2 electrode RL1200 | SI
806205065729 | 10321857 | pCo2 electrode RL1200 | SI
806201428622 | 10315219 | Ca electrode RL1200 | SI
8062478740 | 10323928 | Sample ground temp sensor RL1200 | SI
806204142460 | 10309073 | Sampler Port capillary kit RL1200 (adaptador toma de muestra x 20 un) luer pack | SI
806206650153 | 10324921 | Lamp Assy Co-ox | SI
8062122521 | 10318429 | Air filter x 2 | SI
RapidChem 744/754
806202970609 | 10309048 | Reagent module (Na/K/Cl) para RC 744 | SI
806201331777 | 10309011 | Daily cleaning solution kit | SI
806204473866 | 10309075 | Na Sensor | SI
80624955798 | 10309082 | K sensor | SI
8062661439 | 10308989 | Cl Sensor | SI
806202451792 | 10309037 | Separator sensor | SI
806202283180 | 10309034 | Reference Sensor | SI
806200362008 | 10313301 | Compression plate | SI
806207372033 | 10309472 | Pump tubing kit | SI
806205905239 | 10309095 | Probe wiper 4 | SI
806202377673 | 10309036 | Sampler | NO
806200938058 | 10308995 | Urine diluent | SI
806203279837 | 10318630 | Red Test Dye solution x 50 ml | SI
806205840552 | 10309093 | Reagent Module (Na/K/Li) para RC 754 | SI
806200583772 | 10308981 | Li Sensor | SI
806205610166 | 10309089 | Rc Daily Cleaning Soln Kit | SI
ACCESORIOS COMUNES
80634395 | 10333876 | Printer Paper 248 - 348 - 744- 754 , CTK 50, 100, 500, Advantus (1 Rollo) | SI
ELECTRODOS Y CONSUMIBLES COMUNES 248 -348 -1200
8062476246 | 10324408 | pO2 Electrode 1electrodo | SI
8062476247 | 10317498 | pCO2 Electrode1 electrodo | SI
8062476267 | 10312556 | pH Electrode 1 electrodo + ampolleta con fill solution | SI
8062476266 | 10312557 | Na Electrode 1 electrodo + ampolleta con fill solution | SI
8062476270 | 10327404 | K Electrode 1 electrodo + ampolleta con fill solution | SI
8062476279 | 10330133 | Cl Electrode 1 electrodo + ampolleta con fill solution | SI
8062476268 | 10315922 | Ca Electrode1 electrodo + ampolleta con fill solution | SI
8062476378 | 10320748 | Glucose Biosensor 2 biosensores | SI
8062476379 | 10311032 | Lactate Biosensor 4 biosensores | SI
8062478509 | 10329947 | Reference Electrode Inner Element 1 unidad | SI
8062478498 | 10320458 | Reference Electrode Refill Cassette - Carcasa más Piedra 1 Unidad | SI
8062478822 | 10311081 | Reference Electrode Fill Solution KCL 4M 4 ampolletas x 5 ml. | SI
8062478533 | 10311046 | pH Electrode Fill Solution 3ampolletas x 3 ml. | SI
8062478535 | 10311047 | Na-K-Ca-Cl Electrode Fill Solution 3 amp. x 3 ml. | SI
8062105610 | 10309775 | Desproteinizer universal 10 x 2 ml. | SI
8062478701 | 10311078 | Conditioner Kit 5 x 2 ml | SI
8062477570000 | 10324084 | Clot Removing Kit 1 unidad | NO
8062478498 | 10320458 | Reference Electrode Refill Cassette - Carcasa más Piedra 1 Unidad | SI
CONTROLES COMUNES 238 - 248 -348 - 600 - 800 - 1200
8062108860 | 10309925 | Rapid QC Complete Level 1 pack x 30 amp. | SI
8062108868 | 10309926 | Rapid QC Complete Level 2 pack x 30 amp. | SI
8062108869 | 10309927 | Rapid QC Complete Level 3 pack x 30 amp. | SI
CONTROLES RAPIDLAB 1200
8062120241 | 10310323 | Automatic QC Cartridge kit (1 unidad) | SI
CONTROLES RAPIDCHEM 744 / 754
| 10322876 | Bi Level QC KIT ( Normal y Anormal)- Reemplaza Certain Lyte | SI
Rapidpoint 500
| 10491449 | RP500 MCART LAC 750 TEST | NO
| 10329097 | WASH/WASTE KIT 4 CARTRID | NO
| 10310323 | Automatic QC cartridge | NO
| 10315772 | Printer paper RL 1200 / RP500 | NO
| 10309073 | Sampler Port capillary kit RL1200 & RP500 (adaptador toma de muestra x 20 un) pack | NO
| 10319811 | CART KIT REAGENT RLNG | NO

## Instrumentos
Codigo SAP | SISTEMA | Disponibilidad
10488923 | ADVIA 2120i | SI
11170841 | ADVIA 360 Instrument | SI
11170842 | ADVIA 560 Instrument | SI
11170843 | ADVIA 560 autoloader | SI
11170882 | ADVIA 360/560 External Barcode Reader | SI
10309710 | ADVIA CENTAUR CP | SI
10711433 | ADVIA CENTAUR XPT | SI
10636928 | DIMENSION ExL 200 | SI
10315183 | ADVIA 1800 | SI
10380059 | Immulite 1000 | SI
10373214 | Immulite 2000 XpI | SI
10793839 | Versacell X3 | SI
10323577 | RapidLab 248 | SI
10316271 | Rapidlab 348 | SI
| RL348 EX | SI
10321852 | RapidLab 1265 | SI
| RP 500 | SI
10282970 | DCA Vantage | SI
10494134 | Clinitek Novus | SI
10379677 | Clinitek Status + | SI
10376322 | Clinitek Status Connect | SI
10484765 | Clinitek Advantus | SI
10488536 | PFA 200 | SI
10459329 | BFT II | SI
10712039 | CA-660 | SI
10488583 | CS2000i | SI
10461894 | BCS XP | SI
10461865 | BN Prospec | SI